## VS2000
| TG-43 CONSENSUS | [5 empty] | Varian HDR 192Ir VS2000
Dose calculation for photon-emitting brachytherapy sources with average energy higher than 50 keV: Full Report of the AAPM and ESTRO
Med. Phys. 39 (2012) 2904-2929
| L = | 1.1 |  | cGy/(h U)
Interpolated / extrapolated data are boldface / underlined. Values inside the source are in italics.
| gL(r) |  |  |  | F(r,q) | [9 empty] | QA Along-away
| L = | 0.5 | cm |  |  |  |  | Distance from Active Source Center (cm)
| r (cm) | gL(r) L = 0,35 cm unbounded |  | Theta/deg | 0 | 0.25 | 0.5 | 1 | 2 | 3 | 4 | 5 | 7.5 | 10 | z / cm | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7
| 0 | 0.98607 |  | 0 | 0.59755 | 0.59755 | 0.5729 | 0.5236 | 0.5359 | 0.5679 | 0.5994 | 0.6265 | 0.6875 | 0.734 | 7 | 0.015373 | 0.016025 | 0.016882 | 0.017621 | 0.018151 | 0.018711 | 0.018674 | 0.017715 | 0.016097 | 0.014216 | 0.012351 | 0.010659
| 0.2 | 0.98607 |  | 1 | 0.50783 | 0.50783 | 0.51405 | 0.5265 | 0.5514 | 0.5803 | 0.6119 | 0.6383 | 0.698 | 0.743 | 6 | 0.020388 | 0.02159 | 0.023103 | 0.024253 | 0.025057 | 0.025587 | 0.025277 | 0.023287 | 0.020485 | 0.017571 | 0.014847 | 0.012492
| 0.25 | 0.99101 |  | 2 | 0.52622 | 0.52622 | 0.53295 | 0.5464 | 0.5733 | 0.6055 | 0.635 | 0.6608 | 0.7147 | 0.7568 | 5 | 0.0285 | 0.030878 | 0.033591 | 0.03545 | 0.036549 | 0.036783 | 0.035688 | 0.031506 | 0.026457 | 0.021751 | 0.017812 | 0.014532
| 0.5 | 0.99717 |  | 3 | 0.55645 | 0.55645 | 0.5628 | 0.5755 | 0.6009 | 0.6313 | 0.6583 | 0.6835 | 0.7327 | 0.7707 | 4 | 0.042769 | 0.047891 | 0.053307 | 0.056412 | 0.057434 | 0.056535 | 0.053025 | 0.043736 | 0.034458 | 0.026873 | 0.02109 | 0.016789
| 0.75 | 0.99898 |  | 5 | 0.62567 | 0.62567 | 0.63105 | 0.6418 | 0.6633 | 0.6904 | 0.7116 | 0.7316 | 0.7721 | 0.8036 | 3 | 0.0722 | 0.086306 | 0.09732 | 0.10137 | 0.10126 | 0.094279 | 0.083533 | 0.061632 | 0.044612 | 0.032786 | 0.02467 | 0.018993
| 1 | 1 |  | 7 | 0.69125 | 0.69125 | 0.6951 | 0.7028 | 0.7182 | 0.7387 | 0.7562 | 0.771 | 0.8042 | 0.8288 | 2 | 0.15433 | 0.20437 | 0.22641 | 0.22514 | 0.21188 | 0.17492 | 0.13879 | 0.086373 | 0.056402 | 0.038838 | 0.02795 | 0.020895
| 1.5 | 1.0051 |  | 10 | 0.77015 | 0.77015 | 0.7717 | 0.7748 | 0.781 | 0.7969 | 0.8106 | 0.8212 | 0.8442 | 0.8635 | 1.5 | 0.27324 | 0.3838 | 0.40486 | 0.3778 | 0.33433 | 0.24653 | 0.17896 | 0.10029 | 0.061961 | 0.041438 | 0.029299 | 0.021657
| 2 | 1.0101 |  | 12 | 0.80125 | 0.80125 | 0.8028 | 0.8059 | 0.8121 | 0.8246 | 0.8348 | 0.8431 | 0.8632 | 0.8776 | 1 | 0.62695 | 0.92984 | 0.86367 | 0.7053 | 0.55456 | 0.34403 | 0.22482 | 0.11306 | 0.066671 | 0.043467 | 0.030329 | 0.022229
| 3 | 1.0122 |  | 15 | 0.84237 | 0.84237 | 0.84315 | 0.8447 | 0.8478 | 0.8578 | 0.8647 | 0.8708 | 0.8858 | 0.898 | 0.5 | 3.4202 | 3.8657 | 2.2692 | 1.3703 | 0.89103 | 0.44783 | 0.26507 | 0.12238 | 0.069759 | 0.044852 | 0.031004 | 0.022583
| 4 | 1.013 |  | 20 | 0.88888 | 0.88888 | 0.88925 | 0.89 | 0.8915 | 0.8979 | 0.9024 | 0.9039 | 0.9148 | 0.9236 | 0 | 238116072.079 | 13.9795 | 4.1519 | 1.9243 | 1.1 | 0.49689 | 0.28202 | 0.12596 | 0.070982 | 0.04534 | 0.031251 | 0.022726
| 5 | 1.0106 |  | 25 | 0.92905 | 0.92905 | 0.9259 | 0.9196 | 0.9181 | 0.9224 | 0.9248 | 0.9261 | 0.9342 | 0.938 | -0.5 | 1.7863 | 3.8699 | 2.2734 | 1.3728 | 0.89098 | 0.44814 | 0.26513 | 0.12247 | 0.069778 | 0.044849 | 0.031011 | 0.022593
| 6 | 1.0028 |  | 30 | 0.9445 | 0.9445 | 0.9431 | 0.9403 | 0.9378 | 0.9421 | 0.9442 | 0.9455 | 0.9494 | 0.9531 | -1 | 0.37758 | 0.93027 | 0.86334 | 0.70486 | 0.55555 | 0.34401 | 0.22476 | 0.11311 | 0.066702 | 0.043467 | 0.030327 | 0.022229
| 8 | 0.98209 |  | 35 | 0.95895 | 0.95895 | 0.9582 | 0.9567 | 0.9524 | 0.9553 | 0.9561 | 0.9588 | 0.9609 | 0.9648 | -1.5 | 0.17096 | 0.38076 | 0.40498 | 0.37797 | 0.33447 | 0.24664 | 0.17884 | 0.10027 | 0.061922 | 0.041411 | 0.029278 | 0.02164
| 10 | 0.94928 |  | 40 | 0.96905 | 0.96905 | 0.9681 | 0.9662 | 0.9634 | 0.9673 | 0.9678 | 0.9689 | 0.97 | 0.9714 | -2 | 0.1001 | 0.20253 | 0.22637 | 0.22491 | 0.2121 | 0.17503 | 0.13881 | 0.086311 | 0.056403 | 0.038792 | 0.027933 | 0.020871
|  |  |  | 45 | 0.9755 | 0.9755 | 0.9751 | 0.9743 | 0.9734 | 0.9753 | 0.9746 | 0.9762 | 0.9766 | 0.9784 | -3 | 0.048288 | 0.083568 | 0.096821 | 0.10135 | 0.10108 | 0.094408 | 0.083656 | 0.061693 | 0.044636 | 0.032802 | 0.024652 | 0.018981
|  |  |  | 50 | 0.9934 | 0.9934 | 0.9827 | 0.9818 | 0.9793 | 0.9808 | 0.9809 | 0.9799 | 0.9832 | 0.984 | -4 | 0.029402 | 0.045544 | 0.052863 | 0.056201 | 0.057431 | 0.056415 | 0.053144 | 0.043722 | 0.034429 | 0.026847 | 0.02108 | 0.0168
|  |  |  | 55 | 0.9956 | 0.9956 | 0.9873 | 0.9864 | 0.9857 | 0.9869 | 0.9853 | 0.9845 | 0.9851 | 0.9859 | -5 | 0.020212 | 0.029159 | 0.033113 | 0.035332 | 0.036564 | 0.03681 | 0.035738 | 0.031508 | 0.026458 | 0.021742 | 0.017773 | 0.014533
|  |  |  | 60 | 0.9967 | 0.9967 | 0.9911 | 0.9908 | 0.9882 | 0.9915 | 0.9912 | 0.9896 | 0.9907 | 0.9906 | -6 | 0.014927 | 0.020216 | 0.02261 | 0.024125 | 0.024996 | 0.025609 | 0.025278 | 0.023323 | 0.020487 | 0.017572 | 0.014836 | 0.012472
|  |  |  | 65 | 0.9974 | 0.9974 | 0.994 | 0.9971 | 0.9927 | 0.9938 | 0.9937 | 0.9956 | 0.9948 | 0.995 | -7 | 0.011582 | 0.014869 | 0.016416 | 0.017434 | 0.018071 | 0.01871 | 0.018681 | 0.017721 | 0.016117 | 0.014205 | 0.012345 | 0.010643
|  |  |  | 70 | 0.9979 | 0.9979 | 0.9965 | 0.9969 | 0.9949 | 0.9952 | 0.9955 | 0.9965 | 0.9963 | 0.9964
|  |  |  | 75 | 0.9983 | 0.9983 | 0.9984 | 0.9998 | 0.9968 | 0.9982 | 0.9981 | 0.9982 | 0.9974 | 0.9988
|  |  |  | 80 | 0.999 | 0.999 | 0.9984 | 1.0021 | 0.9979 | 0.9982 | 0.997 | 0.9974 | 0.9991 | 0.9979
|  |  |  | 85 | 0.9992 | 0.9992 | 0.9998 | 1.0006 | 0.9984 | 0.9987 | 0.9989 | 0.9996 | 0.9985 | 0.9999
|  |  |  | 90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  |  | 95 | 0.999 | 0.999 | 1.0007 | 0.9995 | 0.9978 | 0.9986 | 0.9976 | 0.9995 | 0.9992 | 0.9985
|  |  |  | 100 | 0.9988 | 0.9988 | 0.9983 | 1.005 | 0.9974 | 0.999 | 0.9994 | 0.9975 | 0.9987 | 1
|  |  |  | 105 | 0.9988 | 0.9988 | 0.9981 | 0.9997 | 0.9971 | 0.9999 | 0.9982 | 0.998 | 0.9958 | 0.997
|  |  |  | 110 | 0.9989 | 0.9989 | 0.9965 | 0.9994 | 0.9947 | 0.9953 | 0.9949 | 0.9951 | 0.9967 | 0.997
|  |  |  | 115 | 0.9981 | 0.9981 | 0.9944 | 0.9949 | 0.9916 | 0.9938 | 0.9954 | 0.9943 | 0.9936 | 0.9928
|  |  |  | 120 | 0.9966 | 0.9966 | 0.9912 | 0.9955 | 0.9901 | 0.9896 | 0.9895 | 0.9903 | 0.9909 | 0.9934
|  |  |  | 125 | 0.9952 | 0.9952 | 0.9873 | 0.9878 | 0.9842 | 0.9867 | 0.9849 | 0.9852 | 0.9842 | 0.9878
|  |  |  | 130 | 0.9934 | 0.9934 | 0.9812 | 0.9788 | 0.9783 | 0.981 | 0.9806 | 0.9801 | 0.9809 | 0.9826
|  |  |  | 135 | 0.976 | 0.976 | 0.9763 | 0.9769 | 0.9739 | 0.9754 | 0.9762 | 0.9752 | 0.9763 | 0.9769
|  |  |  | 140 | 0.97055 | 0.97055 | 0.9694 | 0.9671 | 0.964 | 0.9674 | 0.9681 | 0.9691 | 0.9702 | 0.9707
|  |  |  | 145 | 0.9603 | 0.9603 | 0.9584 | 0.9546 | 0.9529 | 0.9563 | 0.9582 | 0.9582 | 0.9607 | 0.9633
|  |  |  | 150 | 0.94715 | 0.94715 | 0.9452 | 0.9413 | 0.9395 | 0.944 | 0.944 | 0.9451 | 0.9507 | 0.9537
|  |  |  | 155 | 0.93085 | 0.93085 | 0.9267 | 0.9184 | 0.9186 | 0.9233 | 0.9269 | 0.9305 | 0.9348 | 0.942
|  |  |  | 160 | 0.89242 | 0.89242 | 0.89215 | 0.8916 | 0.8905 | 0.896 | 0.899 | 0.9038 | 0.9143 | 0.9207
|  |  |  | 165 | 0.84475 | 0.84475 | 0.8452 | 0.8461 | 0.8479 | 0.8576 | 0.8649 | 0.8719 | 0.8863 | 0.8992
|  |  |  | 168 | 0.79873 | 0.79873 | 0.80055 | 0.8042 | 0.8115 | 0.8245 | 0.8337 | 0.8446 | 0.8629 | 0.8779
|  |  |  | 170 | 0.75472 | 0.75472 | 0.75825 | 0.7653 | 0.7794 | 0.7935 | 0.8064 | 0.8195 | 0.843 | 0.8611
|  |  |  | 173 | 0.6758 | 0.6758 | 0.6809 | 0.6911 | 0.7115 | 0.7319 | 0.7497 | 0.7674 | 0.799 | 0.8262
|  |  |  | 175 | 0.6075 | 0.6075 | 0.6133 | 0.6249 | 0.6481 | 0.6713 | 0.6945 | 0.7171 | 0.7599 | 0.7929
|  |  |  | 177 | 0.51118 | 0.51118 | 0.51835 | 0.5327 | 0.5614 | 0.5901 | 0.6188 | 0.6475 | 0.7041 | 0.7489
|  |  |  | 178 | 0.51382 | 0.51382 | 0.51894 | 0.52918 | 0.54966 | 0.57014 | 0.59062 | 0.6111 | 0.6623 | 0.7135
|  |  |  | 179 | 0.40249 | 0.40249 | 0.40908 | 0.42226 | 0.44862 | 0.47498 | 0.50134 | 0.5277 | 0.5936 | 0.6595
|  |  |  | 180 | 0.29116 | 0.29116 | 0.29922 | 0.31534 | 0.34758 | 0.37982 | 0.41206 | 0.4443 | 0.5249 | 0.6055