## GI192M11
| TG-43 CONSENSUS | [5 empty] | E&Z BEBIG HDR 192Ir model GI192M11
Dose calculation for photon-emitting brachytherapy sources with average energy higher than 50 keV: Full Report of the AAPM and ESTRO
Med. Phys. 39 (2012) 2904-2929
| L = | 1.11 |  | cGy/(h U)
Interpolated / extrapolated data are boldface / underlined. Values inside the source are in italics.
| gL(r) |  |  |  | F(r,q) | [9 empty] | QA Along-away
| L = | 0.35 | cm |  |  |  |  | Distance from Active Source Center (cm)
| r (cm) | gL( r) |  | Theta/deg | 0 | 0.25 | 0.5 | 1 | 2 | 3 | 4 | 5 | 7.5 | 10 | z / cm | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7
| 0 | 0.9903 |  | 0 | 0.6684 | 0.6684 | 0.6505 | 0.6147 | 0.6211 | 0.644 | 0.6669 | 0.6894 | 0.7335 | 0.769 | 7 | 0.016284 | 0.016613 | 0.016917 | 0.017175 | 0.017447 | 0.01773 | 0.01774 | 0.016968 | 0.015508 | 0.013794 | 0.011989 | 0.010397
| 0.25 | 0.9903 |  | 1 | 0.6709 | 0.6709 | 0.6504 | 0.6094 | 0.6247 | 0.6477 | 0.6725 | 0.6954 | 0.7395 | 0.7725 | 6 | 0.021853 | 0.022403 | 0.022877 | 0.023404 | 0.023859 | 0.024197 | 0.024072 | 0.022381 | 0.019878 | 0.017084 | 0.014489 | 0.012217
| 0.5 | 0.9964 |  | 2 | 0.6676 | 0.6676 | 0.6487 | 0.6109 | 0.6337 | 0.6594 | 0.682 | 0.7048 | 0.7485 | 0.7799 | 5 | 0.030926 | 0.031923 | 0.032846 | 0.033836 | 0.034459 | 0.034808 | 0.034076 | 0.030327 | 0.025736 | 0.021284 | 0.017461 | 0.014284
| 0.75 | 0.9984 |  | 3 | 0.6603 | 0.6603 | 0.6486 | 0.6252 | 0.6466 | 0.6733 | 0.6941 | 0.7147 | 0.7582 | 0.7875 | 4 | 0.046988 | 0.049099 | 0.051175 | 0.052966 | 0.053872 | 0.053676 | 0.050745 | 0.042374 | 0.033683 | 0.026433 | 0.020781 | 0.016508
| 1 | 1 |  | 5 | 0.66825 | 0.66825 | 0.6616 | 0.6483 | 0.6706 | 0.6934 | 0.7127 | 0.7313 | 0.7703 | 0.7963 | 3 | 0.080924 | 0.086253 | 0.091605 | 0.094719 | 0.095308 | 0.090132 | 0.080732 | 0.060289 | 0.043847 | 0.032311 | 0.024349 | 0.018738
| 1.5 | 1.0025 |  | 7 | 0.69095 | 0.69095 | 0.6862 | 0.6767 | 0.6957 | 0.7177 | 0.7358 | 0.7538 | 0.7881 | 0.8129 | 2 | 0.17606 | 0.19469 | 0.20974 | 0.21106 | 0.20162 | 0.1694 | 0.13572 | 0.085053 | 0.055643 | 0.0383 | 0.027612 | 0.020661
| 2 | 1.0035 |  | 10 | 0.72735 | 0.72735 | 0.724 | 0.7173 | 0.736 | 0.7556 | 0.771 | 0.7873 | 0.8159 | 0.8355 | 1.5 | 0.31295 | 0.35441 | 0.37576 | 0.3575 | 0.32112 | 0.24052 | 0.17596 | 0.099223 | 0.061282 | 0.040902 | 0.028969 | 0.021423
| 3 | 1.0052 |  | 12 | 0.7547 | 0.7547 | 0.752 | 0.7466 | 0.7626 | 0.7793 | 0.7919 | 0.8044 | 0.8298 | 0.8497 | 1 | 0.711 | 0.83779 | 0.80594 | 0.6759 | 0.54014 | 0.3387 | 0.22255 | 0.11225 | 0.065966 | 0.042972 | 0.03001 | 0.022005
| 4 | 1.0035 |  | 15 | 0.7959 | 0.7959 | 0.7921 | 0.7845 | 0.7971 | 0.8103 | 0.821 | 0.8317 | 0.8525 | 0.8683 | 0.5 | 3.3128 | 3.4263 | 2.1881 | 1.3478 | 0.88174 | 0.44579 | 0.2634 | 0.12171 | 0.069207 | 0.04437 | 0.030698 | 0.022369
| 5 | 0.999 |  | 20 | 0.84875 | 0.84875 | 0.8455 | 0.839 | 0.8453 | 0.8566 | 0.8652 | 0.873 | 0.8865 | 0.8966 | 0 | 381686141.157 | 15.5001 | 4.2987 | 1.9551 | 1.11 | 0.49733 | 0.28058 | 0.12509 | 0.070278 | 0.044786 | 0.030872 | 0.022452
| 6 | 0.9915 |  | 25 | 0.89125 | 0.89125 | 0.885 | 0.8725 | 0.8818 | 0.8891 | 0.8945 | 0.8982 | 0.9093 | 0.9163 | -0.5 | 1.2812 | 3.4372 | 2.1925 | 1.354 | 0.88383 | 0.44566 | 0.26329 | 0.12179 | 0.069342 | 0.044363 | 0.030683 | 0.022361
| 8 | 0.9682 |  | 30 | 0.9139 | 0.9139 | 0.909 | 0.8992 | 0.9081 | 0.9133 | 0.9171 | 0.9179 | 0.9296 | 0.9324 | -1 | 0.30971 | 0.83305 | 0.80998 | 0.67782 | 0.54122 | 0.33941 | 0.22224 | 0.11209 | 0.066068 | 0.043058 | 0.030036 | 0.021999
| 10 | 0.9352 |  | 35 | 0.9372 | 0.9372 | 0.9327 | 0.9237 | 0.9298 | 0.9342 | 0.9376 | 0.9402 | 0.9473 | 0.9509 | -1.5 | 0.14381 | 0.3435 | 0.37532 | 0.35837 | 0.32178 | 0.24081 | 0.17619 | 0.099098 | 0.061143 | 0.040918 | 0.029005 | 0.021432
|  |  |  | 40 | 0.956 | 0.956 | 0.9506 | 0.9398 | 0.9458 | 0.9492 | 0.9509 | 0.9534 | 0.9578 | 0.9567 | -2 | 0.085072 | 0.18371 | 0.20827 | 0.21103 | 0.20184 | 0.16967 | 0.13592 | 0.08511 | 0.055651 | 0.03827 | 0.027598 | 0.020679
|  |  |  | 45 | 0.96495 | 0.96495 | 0.9616 | 0.9549 | 0.9591 | 0.9631 | 0.9646 | 0.9662 | 0.9682 | 0.9699 | -3 | 0.041779 | 0.076891 | 0.088993 | 0.094062 | 0.094982 | 0.090225 | 0.08084 | 0.060297 | 0.043912 | 0.03231 | 0.024311 | 0.018691
|  |  |  | 50 | 0.9743 | 0.9743 | 0.972 | 0.9674 | 0.9703 | 0.973 | 0.974 | 0.9765 | 0.9784 | 0.9775 | -4 | 0.025706 | 0.04191 | 0.048775 | 0.051983 | 0.053515 | 0.053508 | 0.05079 | 0.042422 | 0.033654 | 0.026438 | 0.020794 | 0.016495
|  |  |  | 55 | 0.9862 | 0.9862 | 0.9799 | 0.9716 | 0.9772 | 0.9796 | 0.9794 | 0.9796 | 0.9837 | 0.9853 | -5 | 0.017818 | 0.026868 | 0.030662 | 0.032687 | 0.033971 | 0.034628 | 0.034002 | 0.0304 | 0.025764 | 0.02126 | 0.01746 | 0.014294
|  |  |  | 60 | 0.99 | 0.99 | 0.985 | 0.9777 | 0.9848 | 0.9879 | 0.989 | 0.9878 | 0.9894 | 0.9884 | -6 | 0.013276 | 0.018638 | 0.021071 | 0.022445 | 0.023259 | 0.024026 | 0.023973 | 0.022385 | 0.019861 | 0.017107 | 0.014476 | 0.01222
|  |  |  | 65 | 0.9929 | 0.9929 | 0.9904 | 0.9862 | 0.99 | 0.9924 | 0.9927 | 0.9924 | 0.9956 | 0.9936 | -7 | 0.01038 | 0.013716 | 0.015365 | 0.016325 | 0.016903 | 0.017553 | 0.017651 | 0.016935 | 0.015531 | 0.013772 | 0.012026 | 0.010392
|  |  |  | 70 | 0.9958 | 0.9958 | 0.9954 | 0.9921 | 0.9932 | 0.9967 | 0.996 | 0.9954 | 0.9968 | 0.9964
|  |  |  | 75 | 0.9978 | 0.9978 | 0.9971 | 0.9961 | 0.996 | 0.9973 | 0.9975 | 0.9975 | 0.9998 | 0.9983
|  |  |  | 80 | 0.9996 | 0.9996 | 1.001 | 0.9949 | 0.9989 | 1.0001 | 0.9984 | 0.9988 | 1.0008 | 1.0006
|  |  |  | 85 | 0.9998 | 0.9998 | 1.0017 | 0.9961 | 0.9988 | 0.999 | 1.0016 | 1.0011 | 1.0022 | 1.0021
|  |  |  | 90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  |  | 95 | 0.9997 | 0.9997 | 1.0006 | 0.9952 | 0.9991 | 1.0017 | 1.0022 | 1.0005 | 1.0018 | 1.0023
|  |  |  | 100 | 0.9991 | 0.9991 | 0.999 | 0.9967 | 0.998 | 1.0004 | 1.0023 | 1.0012 | 1.0001 | 1.0017
|  |  |  | 105 | 0.9986 | 0.9986 | 0.9983 | 0.9957 | 0.9956 | 0.9983 | 0.9985 | 0.9987 | 1.0012 | 1
|  |  |  | 110 | 0.996 | 0.996 | 0.9947 | 0.9911 | 0.9935 | 0.9943 | 0.9928 | 0.9942 | 0.9953 | 0.9931
|  |  |  | 115 | 0.9939 | 0.9939 | 0.9918 | 0.988 | 0.9884 | 0.991 | 0.9937 | 0.993 | 0.9925 | 0.9924
|  |  |  | 120 | 0.9896 | 0.9896 | 0.989 | 0.9827 | 0.9845 | 0.9848 | 0.9868 | 0.9874 | 0.9886 | 0.9882
|  |  |  | 125 | 0.9856 | 0.9856 | 0.9815 | 0.9766 | 0.9791 | 0.9811 | 0.9814 | 0.9823 | 0.9846 | 0.9859
|  |  |  | 130 | 0.97225 | 0.97225 | 0.9708 | 0.9679 | 0.9706 | 0.9739 | 0.9738 | 0.9759 | 0.9782 | 0.9782
|  |  |  | 135 | 0.96595 | 0.96595 | 0.9631 | 0.9574 | 0.9602 | 0.9646 | 0.965 | 0.9655 | 0.967 | 0.9695
|  |  |  | 140 | 0.9546 | 0.9546 | 0.9503 | 0.9417 | 0.9477 | 0.9507 | 0.9536 | 0.9561 | 0.9588 | 0.9612
|  |  |  | 145 | 0.93785 | 0.93785 | 0.9344 | 0.9275 | 0.9308 | 0.936 | 0.9379 | 0.9403 | 0.9455 | 0.9487
|  |  |  | 150 | 0.91885 | 0.91885 | 0.9143 | 0.9052 | 0.9102 | 0.9164 | 0.9188 | 0.9217 | 0.9306 | 0.9356
|  |  |  | 155 | 0.891 | 0.891 | 0.8863 | 0.8769 | 0.8821 | 0.8894 | 0.8944 | 0.8982 | 0.9084 | 0.9137
|  |  |  | 160 | 0.83272 | 0.83272 | 0.83455 | 0.8382 | 0.8455 | 0.8541 | 0.8621 | 0.8701 | 0.8834 | 0.8928
|  |  |  | 165 | 0.7743 | 0.7743 | 0.7771 | 0.7827 | 0.7939 | 0.8068 | 0.8175 | 0.8265 | 0.848 | 0.8626
|  |  |  | 168 | 0.72542 | 0.72542 | 0.72925 | 0.7369 | 0.7522 | 0.7693 | 0.7828 | 0.7967 | 0.8213 | 0.8416
|  |  |  | 170 | 0.68465 | 0.68465 | 0.6895 | 0.6992 | 0.7186 | 0.738 | 0.7542 | 0.7691 | 0.7997 | 0.8221
|  |  |  | 173 | 0.61445 | 0.61445 | 0.6203 | 0.632 | 0.6554 | 0.6788 | 0.7004 | 0.7191 | 0.7595 | 0.7882
|  |  |  | 175 | 0.55383 | 0.55383 | 0.56035 | 0.5734 | 0.5995 | 0.6256 | 0.6517 | 0.674 | 0.7231 | 0.7587
|  |  |  | 177 | 0.45817 | 0.45817 | 0.46595 | 0.4815 | 0.5126 | 0.5437 | 0.5748 | 0.6059 | 0.6684 | 0.7131
|  |  |  | 178 | 0.46127 | 0.46127 | 0.46674 | 0.47768 | 0.49956 | 0.52144 | 0.54332 | 0.5652 | 0.6199 | 0.6746
|  |  |  | 179 | 0.33042 | 0.33042 | 0.33764 | 0.35208 | 0.38096 | 0.40984 | 0.43872 | 0.4676 | 0.5398 | 0.612
|  |  |  | 180 | 0.24349 | 0.24349 | 0.25158 | 0.26776 | 0.30012 | 0.33248 | 0.36484 | 0.3972 | 0.4781 | 0.559