## General Information
General Information
Source Information
Source name: | HDR Varian VariSource (classic model)
Radionuclide: | Ir-192
Model Number: | VariSource (classic model)
Distributed by: | Varian Oncology Systems
Manufactured by:
References
1) | "Monte Carlo dosimetry of the VariSource high dose rate192Ir source"
Med Phys 25(1998), 415-423
R. Wang, R. S. Sloboda
Other References
1) | "A Monte Carlo investigation of the dosimetric characteristics of the
VariSource 192Ir high dose rate brachytherapy source"
Med Phys 26(1999), 1498-1502
P. Karaiskos, A. Angelopoulos, P.Baras, L. Sakelliou,
P. Sandilos, K. Dardoufas, L. Vlachos
2) | "Dosimetric characteristics of a new high-intensity 192Ir source
for remote afterloading"
Med Phys 24(1997), 2008-2013
A. S. Meigooni, M. T. Kleiman, J. L. Johnson,
D. Mazloomdoost, G. S. Ibbott
3) | "Dose characterization in the near-source region for two high
dose rate brachytherapy sources"
Med Phys 29(2002), 1678-1686
R. Wang, X. A. Li
4) | "Monte Carlo dosimetry of a new 192Ir high
dose rate brachytherapy source"
Med Phys 27(2000), 2521-2527
A. Angelopoulos, P.Baras, L. Sakelliou,
P. Karaiskos, P. Sandilos
5) | "Fitted dosimetric parameters of high dose-rate 192Ir
sources according to the AAPM TG43 formalism"
Med Phys 28(2001), 654-660
F. Lliso, J. Pérez-Calatayud, V. Carmona, F. Ballester,
J.L. Lluch, M. A. Serrano, Y. Limami, E. Casal
6) | "Erratum: "Fitted dosimetric parameters of high dose-rate 192Ir sources
according to the AAPM TG43 formalism" [med. Phys. 28, 654-660 (2001)] "
Med Phys 28(2001), 1964
F. Lliso, J. Pérez-Calatayud, V. Carmona, F. Ballester,
J.L. Lluch, M. A. Serrano, Y. Limami, E. Casal
Calculation details
Method: | Monte Carlo
Code: | EGS4
Version:
Phantom material: | Water
Phantom size: | 30 cm in diameter sphere
Cutoff energy:
Cross-section:
Electron tracking: | No
Revised
2006-01-30

## Source design and materials
Source design and materials |  | 192Ir -HDR classic VariSource
Coordinate sistem is located at the source center.
Image has been obtained from Karaiskos et al MPH (1999)
z
y

## Geometry Function
Geometry Function | 192Ir -HDR classic VariSource
L=1 cm is the active length of the source

## Dose Rate Constant
Dose-Rate Constant |  |  | 192Ir -HDR classic VariSource
L = | 1,044 (0,2%) | cGy/(hU)

## D(y,z)
Absolute dose rate table: D(y,z) (cGy/h U)
192Ir -HDR classic VariSource
Note: Positive z axis is directed towards the source tip (in the original article z axis direction is inverted)
y (cm)
z (cm) | 0 | 0.1 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 5 | 6 | 7 | 10
10 | 0.0057 | 0.0065 | 0.0066 | 0.0071 | 0.0074 | 0.0077 | 0.008 | 0.0082 | 0.0082 | 0.0081 | 0.0077 | 0.0071 | 0.0065 | 0.0058 | 0.0038
7 | 0.0129 | 0.0136 | 0.0142 | 0.0155 | 0.0166 | 0.0173 | 0.018 | 0.0181 | 0.0177 | 0.0172 | 0.0155 | 0.0136 | 0.0116 | 0.0101 | 0.0059
6 | 0.0175 | 0.018 | 0.0193 | 0.0217 | 0.0234 | 0.0243 | 0.025 | 0.0248 | 0.024 | 0.0228 | 0.0199 | 0.017 | 0.0142 | 0.0118 | 0.0067
5 | 0.0241 | 0.0253 | 0.0282 | 0.032 | 0.0343 | 0.0358 | 0.0364 | 0.0353 | 0.0334 | 0.031 | 0.026 | 0.0212 | 0.0172 | 0.0139 | 0.0075
4 | 0.0347 | 0.0372 | 0.044 | 0.0522 | 0.0557 | 0.057 | 0.0563 | 0.0528 | 0.0482 | 0.0433 | 0.034 | 0.0264 | 0.0206 | 0.0162 | 0.0083
3 | 0.0558 | 0.0647 | 0.0825 | 0.0967 | 0.1018 | 0.1021 | 0.0949 | 0.0837 | 0.072 | 0.0614 | 0.0441 | 0.0323 | 0.0241 | 0.0184 | 0.0089
2.5 | 0.078 | 0.0969 | 0.1237 | 0.1444 | 0.1489 | 0.1452 | 0.128 | 0.1076 | 0.0888 | 0.073 | 0.0499 | 0.0354 | 0.0258 | 0.0194 | 0.0092
2 | 0.1238 | 0.1614 | 0.2073 | 0.234 | 0.2321 | 0.2175 | 0.177 | 0.1393 | 0.1091 | 0.086 | 0.0559 | 0.0383 | 0.0274 | 0.0203 | 0.0095
1.5 | 0.2292 | 0.3173 | 0.4144 | 0.4337 | 0.3976 | 0.3463 | 0.2495 | 0.1794 | 0.1319 | 0.0998 | 0.0614 | 0.0409 | 0.0287 | 0.0211 | 0.0097
1 | 0.6676 | 0.9841 | 1.142 | 0.9742 | 0.7502 | 0.5706 | 0.3449 | 0.2239 | 0.1547 | 0.1123 | 0.0661 | 0.0429 | 0.0298 | 0.0217 | 0.0098
0.75 | 1.806 | 2.5951 | 2.3897 | 1.5627 | 1.0378 | 0.7225 | 0.3962 | 0.2449 | 0.1646 | 0.1175 | 0.0679 | 0.0437 | 0.0302 | 0.0219 | 0.0099
0.5 |  | 15.435 | 5.7097 | 2.4301 | 1.3746 | 0.8779 | 0.4418 | 0.2622 | 0.1724 | 0.1216 | 0.0692 | 0.0443 | 0.0304 | 0.022 | 0.0099
0.25 |  | 28.207 | 8.9179 | 3.2147 | 1.6503 | 0.9981 | 0.4733 | 0.2735 | 0.1774 | 0.1242 | 0.0701 | 0.0446 | 0.0306 | 0.0222 | 0.0099
0.1 |  | 29.225 | 9.5971 | 3.4506 | 1.7406 | 1.0347 | 0.4827 | 0.2768 | 0.1787 | 0.1251 | 0.07 | 0.0448 | 0.0307 | 0.0221 | 0.01
0 |  | 29.418 | 9.7037 | 3.4958 | 1.7563 | 1.0432 | 0.4842 | 0.2777 | 0.1793 | 0.1247 | 0.0705 | 0.0447 | 0.0307 | 0.0221 | 0.01
-0.1 |  | 29.268 | 9.5966 | 3.4502 | 1.7394 | 1.0342 | 0.4833 | 0.2768 | 0.1788 | 0.1245 | 0.0701 | 0.0448 | 0.0307 | 0.0222 | 0.01
-0.25 |  | 28.252 | 8.909 | 3.2076 | 1.6511 | 0.9978 | 0.4725 | 0.2734 | 0.177 | 0.1239 | 0.0701 | 0.0445 | 0.0306 | 0.022 | 0.0099
-0.5 |  | 15.424 | 5.7044 | 2.4362 | 1.373 | 0.8778 | 0.4411 | 0.2619 | 0.1721 | 0.1214 | 0.0692 | 0.0443 | 0.0304 | 0.022 | 0.0099
-0.75 |  | 2.6069 | 2.3937 | 1.563 | 1.0377 | 0.7228 | 0.396 | 0.2446 | 0.1644 | 0.1175 | 0.0677 | 0.0437 | 0.0302 | 0.0219 | 0.0099
-1 | 0.6072 | 0.9784 | 1.1398 | 0.9753 | 0.7501 | 0.5711 | 0.3452 | 0.224 | 0.1549 | 0.1124 | 0.066 | 0.0429 | 0.03 | 0.0217 | 0.0098
-1.5 | 0.2218 | 0.3094 | 0.4123 | 0.4345 | 0.3979 | 0.3458 | 0.2494 | 0.1794 | 0.132 | 0.0998 | 0.0614 | 0.0407 | 0.0288 | 0.0211 | 0.0097
-2 | 0.1227 | 0.1524 | 0.2059 | 0.2346 | 0.2321 | 0.2176 | 0.1774 | 0.1394 | 0.1091 | 0.0861 | 0.0559 | 0.0382 | 0.0274 | 0.0203 | 0.0095
-2.5 | 0.0789 | 0.0926 | 0.1232 | 0.1442 | 0.1487 | 0.145 | 0.1281 | 0.1077 | 0.0889 | 0.073 | 0.05 | 0.0353 | 0.0258 | 0.0194 | 0.0092
-3 | 0.057 | 0.0633 | 0.0815 | 0.0969 | 0.1021 | 0.1021 | 0.095 | 0.0838 | 0.0721 | 0.0614 | 0.0443 | 0.0323 | 0.0241 | 0.0184 | 0.0089
-4 | 0.0338 | 0.0349 | 0.0434 | 0.0514 | 0.0554 | 0.0569 | 0.0563 | 0.0529 | 0.0482 | 0.0433 | 0.034 | 0.0264 | 0.0206 | 0.0162 | 0.0083
-5 | 0.0239 | 0.0245 | 0.0274 | 0.0317 | 0.0343 | 0.0357 | 0.0364 | 0.0353 | 0.0334 | 0.0311 | 0.0259 | 0.0212 | 0.0172 | 0.0139 | 0.0075
-6 | 0.0162 | 0.0172 | 0.0192 | 0.0215 | 0.0231 | 0.0242 | 0.025 | 0.0247 | 0.0239 | 0.0228 | 0.02 | 0.0169 | 0.0142 | 0.0118 | 0.0067
-7 | 0.013 | 0.0129 | 0.0138 | 0.0154 | 0.0165 | 0.0173 | 0.018 | 0.018 | 0.0178 | 0.0172 | 0.0155 | 0.0136 | 0.0117 | 0.01 | 0.0059
-10 | 0.0061 | 0.0063 | 0.0066 | 0.0071 | 0.0074 | 0.0076 | 0.008 | 0.0081 | 0.0081 | 0.0081 | 0.0077 | 0.0071 | 0.0065 | 0.0058 | 0.0038

## Anisotropy Function Table
| Anisotropy Function Table |  |  |  | 192Ir -HDR classic VariSource
| F(r,q) |  | Note: q=0 indicates the source tip (in the original article q=0 indicates the source cable)
r (cm)
q (º) | 0.25 | 0.5 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 10
0 |  |  | 0.445 | 0.41 | 0.431 | 0.482 | 0.531 | 0.566 | 0.577 | 0.571 | 180
1 |  |  | 0.448 | 0.439 | 0.474 | 0.488 | 0.562 | 0.586 | 0.622 | 0.659 | 179
2 |  |  | 0.469 | 0.472 | 0.503 | 0.551 | 0.583 | 0.603 | 0.637 | 0.669 | 178
3 |  |  | 0.54 | 0.545 | 0.558 | 0.588 | 0.628 | 0.651 | 0.67 | 0.719 | 177
5 |  | 0.943 | 0.637 | 0.624 | 0.651 | 0.671 | 0.688 | 0.71 | 0.724 | 0.758 | 175
7 |  | 0.944 | 0.716 | 0.694 | 0.714 | 0.733 | 0.749 | 0.765 | 0.774 | 0.796 | 173
10 | 0.974 | 0.943 | 0.787 | 0.771 | 0.779 | 0.788 | 0.798 | 0.812 | 0.822 | 0.837 | 170
12 | 0.974 | 0.948 | 0.822 | 0.804 | 0.813 | 0.822 | 0.831 | 0.836 | 0.847 | 0.863 | 168
15 | 0.973 | 0.953 | 0.863 | 0.841 | 0.849 | 0.851 | 0.864 | 0.868 | 0.875 | 0.887 | 165
20 | 0.974 | 0.963 | 0.904 | 0.887 | 0.893 | 0.892 | 0.898 | 0.903 | 0.908 | 0.917 | 160
25 | 0.979 | 0.973 | 0.929 | 0.917 | 0.92 | 0.921 | 0.922 | 0.927 | 0.931 | 0.936 | 155
30 | 0.984 | 0.98 | 0.947 | 0.938 | 0.94 | 0.938 | 0.941 | 0.943 | 0.947 | 0.95 | 150
35 | 0.988 | 0.984 | 0.961 | 0.954 | 0.954 | 0.951 | 0.956 | 0.957 | 0.96 | 0.954 | 145
45 | 0.993 | 0.985 | 0.977 | 0.973 | 0.974 | 0.971 | 0.974 | 0.975 | 0.977 | 0.986 | 135
50 | 0.995 | 0.988 | 0.983 | 0.98 | 0.981 | 0.977 | 0.979 | 0.981 | 0.983 | 0.988 | 130
60 | 0.998 | 0.995 | 0.991 | 0.988 | 0.99 | 0.987 | 0.988 | 0.99 | 0.993 | 0.994 | 120
75 | 1 | 0.999 | 0.998 | 0.995 | 0.999 | 0.995 | 0.997 | 0.997 | 0.999 | 0.998 | 105
90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 90
105 | 1 | 0.999 | 0.998 | 0.994 | 0.999 | 0.994 | 0.994 | 0.997 | 1 | 0.998 | 75
120 | 0.998 | 0.995 | 0.991 | 0.989 | 0.991 | 0.988 | 0.989 | 0.99 | 0.992 | 0.995 | 60
130 | 0.996 | 0.992 | 0.983 | 0.98 | 0.981 | 0.978 | 0.981 | 0.982 | 0.983 | 0.981 | 50
135 | 0.994 | 0.99 | 0.975 | 0.973 | 0.975 | 0.971 | 0.974 | 0.975 | 0.977 | 0.974 | 45
145 | 0.989 | 0.981 | 0.96 | 0.952 | 0.955 | 0.952 | 0.955 | 0.957 | 0.961 | 0.96 | 35
150 | 0.985 | 0.973 | 0.946 | 0.938 | 0.94 | 0.939 | 0.942 | 0.944 | 0.949 | 0.952 | 30
155 | 0.981 | 0.964 | 0.93 | 0.917 | 0.92 | 0.92 | 0.923 | 0.927 | 0.931 | 0.936 | 25
160 | 0.976 | 0.959 | 0.905 | 0.887 | 0.892 | 0.892 | 0.899 | 0.902 | 0.908 | 0.914 | 20
165 | 0.974 | 0.954 | 0.861 | 0.843 | 0.851 | 0.852 | 0.862 | 0.869 | 0.874 | 0.885 | 15
168 | 0.975 | 0.948 | 0.824 | 0.803 | 0.812 | 0.819 | 0.828 | 0.84 | 0.847 | 0.86 | 12
170 | 0.973 | 0.943 | 0.789 | 0.765 | 0.78 | 0.785 | 0.8 | 0.812 | 0.822 | 0.842 | 10
173 |  | 0.943 | 0.712 | 0.69 | 0.713 | 0.722 | 0.742 | 0.755 | 0.771 | 0.793 | 7
175 |  | 0.942 | 0.635 | 0.617 | 0.645 | 0.657 | 0.681 | 0.709 | 0.72 | 0.764 | 5
177 |  |  | 0.51 | 0.512 | 0.544 | 0.581 | 0.614 | 0.645 | 0.669 | 0.709 | 3
178 |  |  | 0.444 | 0.456 | 0.494 | 0.516 | 0.558 | 0.594 | 0.62 | 0.677 | 2
179 |  |  | 0.407 | 0.405 | 0.454 | 0.481 | 0.525 | 0.555 | 0.575 | 0.633 | 1
180 |  |  | 0.405 | 0.406 | 0.441 | 0.469 | 0.528 | 0.524 | 0.582 | 0.606 | 0

## Radial Dose Function
Radial-Dose Function Table |  |  | 192Ir -HDR classic VariSource
| r (cm) | g(r)
| 0.1 | 0.952
| 0.2 | 0.967
| 0.3 | 0.976
| 0.5 | 0.989
| 1 | 1
| 1.5 | 1.005
| 2 | 1.007
| 2.5 | 1.006
| 3 | 1.006
| 4 | 1.003
| 5 | 0.998
| 6 | 0.984
| 7 | 0.967
| 8 | 0.947
| 9 | 0.92
| 10 | 0.885
| 11 | 0.85
| 12 | 0.807
| 13 | 0.757
| 14 | 0.699