## Co60-A86
| TG-43 CONSENSUS | [5 empty] | E&Z BEBIG 60Co HDR Co0-A86
Dose calculation for photon-emitting brachytherapy sources with average energy higher than 50 keV: Full Report of the AAPM and ESTRO
Med. Phys. 39 (2012) 2904-2929
| L = | 1.092 |  | cGy/(h U)
Interpolated / extrapolated data are boldface / underlined. Values inside the source are in italics.
| gL( r) |  |  |  | F(r,q) | [13 empty] | QA Along-away
| L = | 0.35 | cm |  |  |  |  | Distance from Active Source Center (cm)
| r (cm) | gL( r) |  | Theta/deg | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 10 | z / cm | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7
| 0 | 0.83 |  | 0 | 0.9844 | 0.9844 | 0.96504 | 0.94567 | 0.92293 | 0.94482 | 0.94723 | 0.94527 | 0.94371 | 0.94523 | 0.9448 | 0.94697 | 0.94744 | 0.952 | 7 | 0.01923 | 0.019414 | 0.019494 | 0.01941 | 0.019277 | 0.018893 | 0.018323 | 0.01675 | 0.014877 | 0.012952 | 0.011154 | 0.009551
| 0.1 | 0.83 |  | 1 | 0.98118 | 0.98118 | 0.97087 | 0.96057 | 0.95237 | 0.94741 | 0.94742 | 0.9464 | 0.94609 | 0.94829 | 0.94907 | 0.95122 | 0.95225 | 0.95478 | 6 | 0.026627 | 0.026948 | 0.027013 | 0.026827 | 0.026599 | 0.025859 | 0.024793 | 0.022027 | 0.018935 | 0.015984 | 0.013377 | 0.011176
| 0.15 | 0.961 |  | 2 | 0.97737 | 0.97737 | 0.97051 | 0.96365 | 0.96291 | 0.95208 | 0.95169 | 0.95173 | 0.95256 | 0.95492 | 0.95641 | 0.95719 | 0.95863 | 0.95938 | 5 | 0.039085 | 0.039519 | 0.039493 | 0.039144 | 0.038656 | 0.037068 | 0.034916 | 0.029715 | 0.024431 | 0.019805 | 0.016019 | 0.013003
| 0.2 | 1.037 |  | 3 | 0.96296 | 0.96296 | 0.96374 | 0.96452 | 0.95761 | 0.95584 | 0.95718 | 0.95765 | 0.95776 | 0.95877 | 0.96042 | 0.96078 | 0.96194 | 0.96372 | 4 | 0.062057 | 0.062885 | 0.062554 | 0.061701 | 0.060372 | 0.05653 | 0.051643 | 0.041106 | 0.031793 | 0.024494 | 0.01902 | 0.014969
| 0.25 | 1.072 |  | 4 | 0.95072 | 0.95072 | 0.96063 | 0.97053 | 0.96121 | 0.95914 | 0.95966 | 0.96102 | 0.96195 | 0.96248 | 0.96405 | 0.96503 | 0.96644 | 0.96829 | 3 | 0.11253 | 0.11378 | 0.11227 | 0.1093 | 0.10496 | 0.093614 | 0.080928 | 0.057969 | 0.041228 | 0.02986 | 0.0222 | 0.016931
| 0.3 | 1.077 |  | 5 | 0.93848 | 0.93848 | 0.95751 | 0.97653 | 0.9648 | 0.96238 | 0.96198 | 0.96315 | 0.9645 | 0.96498 | 0.96597 | 0.96722 | 0.9681 | 0.9697 | 2 | 0.25901 | 0.25989 | 0.25102 | 0.23521 | 0.21507 | 0.17184 | 0.13363 | 0.081206 | 0.052025 | 0.035321 | 0.025171 | 0.018651
| 0.35 | 1.066 |  | 6 | 0.9391 | 0.9391 | 0.95815 | 0.97721 | 0.9664 | 0.96402 | 0.96395 | 0.96401 | 0.965 | 0.96515 | 0.96601 | 0.96744 | 0.96774 | 0.96962 | 1.5 | 0.46583 | 0.46481 | 0.43425 | 0.38715 | 0.33491 | 0.24059 | 0.17211 | 0.094248 | 0.057237 | 0.037708 | 0.026392 | 0.019332
| 0.4 | 1.05 |  | 8 | 0.94377 | 0.94377 | 0.96161 | 0.97945 | 0.97131 | 0.96742 | 0.96724 | 0.96739 | 0.96806 | 0.96779 | 0.96884 | 0.96976 | 0.9706 | 0.97161 | 1 | 1.0502 | 1.0446 | 0.88775 | 0.70699 | 0.54906 | 0.33489 | 0.21603 | 0.10639 | 0.061596 | 0.039607 | 0.027337 | 0.01985
| 0.45 | 1.037 |  | 10 | 0.95188 | 0.95188 | 0.96723 | 0.98259 | 0.97795 | 0.97109 | 0.97129 | 0.97119 | 0.9716 | 0.97227 | 0.97254 | 0.97294 | 0.97418 | 0.97507 | 0.5 | 4.9882 | 3.8718 | 2.2789 | 1.3656 | 0.87849 | 0.43603 | 0.25469 | 0.11524 | 0.064537 | 0.040845 | 0.02794 | 0.020176
| 0.5 | 1.028 |  | 15 | 0.97001 | 0.97001 | 0.97568 | 0.98134 | 0.98065 | 0.97894 | 0.97876 | 0.97893 | 0.9792 | 0.97914 | 0.97974 | 0.9804 | 0.98037 | 0.98118 | 0 | 463503843.1105 | 16.5067 | 4.3631 | 1.9477 | 1.092 | 0.48414 | 0.27066 | 0.11851 | 0.06559 | 0.041282 | 0.028151 | 0.02029
| 0.6 | 1.019 |  | 20 | 0.73357 | 0.73357 | 0.98226 | 0.98913 | 0.99012 | 0.98466 | 0.98495 | 0.98479 | 0.98508 | 0.98516 | 0.98535 | 0.9856 | 0.98603 | 0.98612 | -0.5 | 5.0374 | 3.8535 | 2.2747 | 1.3649 | 0.8784 | 0.43605 | 0.25473 | 0.11525 | 0.064563 | 0.040857 | 0.027946 | 0.02018
| 0.65 | 1.018 |  | 25 | 0.78784 | 0.78784 | 1.0007 | 0.99483 | 0.9894 | 0.98862 | 0.98879 | 0.98893 | 0.9893 | 0.98945 | 0.98977 | 0.98971 | 0.9899 | 0.98987 | -1 | 1.0591 | 1.0318 | 0.886 | 0.7051 | 0.5488 | 0.33495 | 0.21603 | 0.10639 | 0.06162 | 0.039625 | 0.027349 | 0.019858
| 0.75 | 1.011 |  | 30 | 0.82113 | 0.82113 | 0.9943 | 0.99416 | 0.99108 | 0.99293 | 0.99288 | 0.99258 | 0.99252 | 0.99238 | 0.9927 | 0.99277 | 0.99279 | 0.99241 | -1.5 | 0.44808 | 0.4557 | 0.43161 | 0.38635 | 0.33441 | 0.24054 | 0.17212 | 0.094252 | 0.057241 | 0.037713 | 0.026399 | 0.01934
| 1 | 1 |  | 40 | 0.89384 | 0.89384 | 1.0023 | 0.99882 | 0.99775 | 0.99602 | 0.99615 | 0.99606 | 0.99626 | 0.99576 | 0.99595 | 0.99567 | 0.9961 | 0.99607 | -2 | 0.24791 | 0.25296 | 0.24826 | 0.23421 | 0.21454 | 0.17163 | 0.13357 | 0.081196 | 0.052026 | 0.03532 | 0.025173 | 0.018654
| 1.5 | 0.992 |  | 50 | 0.94204 | 0.94204 | 1.0021 | 0.99816 | 0.99644 | 0.99822 | 0.99823 | 0.99811 | 0.99821 | 0.99803 | 0.99802 | 0.9983 | 0.99822 | 0.99853 | -3 | 0.10793 | 0.10979 | 0.11026 | 0.10809 | 0.10432 | 0.093371 | 0.080795 | 0.057935 | 0.041216 | 0.029855 | 0.022195 | 0.016926
| 2 | 0.984 |  | 60 | 0.97369 | 0.97369 | 0.99823 | 1.0025 | 0.99854 | 0.99909 | 0.99928 | 0.99875 | 0.99867 | 0.99831 | 0.99843 | 0.99877 | 0.99903 | 0.9989 | -4 | 0.059746 | 0.06065 | 0.061008 | 0.060784 | 0.059718 | 0.056296 | 0.051497 | 0.04106 | 0.031772 | 0.02448 | 0.01901 | 0.014962
| 3 | 0.968 |  | 70 | 0.99118 | 0.99118 | 0.9997 | 0.99833 | 1 | 0.99989 | 1 | 0.99957 | 0.99973 | 0.99941 | 0.99951 | 0.9996 | 0.99991 | 0.99981 | -5 | 0.0377 | 0.038165 | 0.038354 | 0.038417 | 0.038125 | 0.036791 | 0.034781 | 0.029651 | 0.024409 | 0.019793 | 0.016016 | 0.012998
| 4 | 0.952 |  | 80 | 0.99978 | 0.99978 | 0.99876 | 0.99987 | 1.0022 | 1.0002 | 1.0002 | 0.99997 | 0.99972 | 0.99956 | 0.99958 | 0.99954 | 0.99968 | 0.99958 | -6 | 0.025775 | 0.026042 | 0.02615 | 0.026224 | 0.02617 | 0.025586 | 0.02465 | 0.021966 | 0.018908 | 0.015972 | 0.01337 | 0.01117
| 5 | 0.936 |  | 90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | -7 | 0.018553 | 0.018772 | 0.018869 | 0.018917 | 0.018909 | 0.018667 | 0.018171 | 0.016696 | 0.01485 | 0.012937 | 0.011145 | 0.0095431
| 6 | 0.919 |  | 100 | 1.0001 | 1.0001 | 0.99741 | 1.0008 | 1.0002 | 1.0004 | 1.0005 | 1.0001 | 1.0003 | 1.0001 | 1 | 1 | 1.0002 | 1
| 8 | 0.884 |  | 110 | 0.99486 | 0.99486 | 1.002 | 1.0002 | 0.99835 | 0.99992 | 0.99991 | 0.9996 | 0.99986 | 0.99935 | 0.9995 | 0.99968 | 0.99949 | 0.99962
| 10 | 0.849 |  | 120 | 0.98221 | 0.98221 | 0.99879 | 1.0022 | 0.99917 | 0.99921 | 0.99925 | 0.99883 | 0.99859 | 0.99842 | 0.99827 | 0.99832 | 0.99859 | 0.99839
|  |  |  | 130 | 0.95706 | 0.95706 | 1.002 | 0.99454 | 0.99572 | 0.99867 | 0.99859 | 0.99793 | 0.9978 | 0.99757 | 0.99747 | 0.99764 | 0.99814 | 0.99773
|  |  |  | 140 | 0.91303 | 0.91303 | 1.0057 | 0.99724 | 0.99249 | 0.99571 | 0.99544 | 0.99521 | 0.99546 | 0.99512 | 0.99515 | 0.99512 | 0.99535 | 0.99519
|  |  |  | 150 | 0.84272 | 0.84272 | 0.99583 | 0.9905 | 0.98812 | 0.99104 | 0.99074 | 0.99041 | 0.99045 | 0.99024 | 0.99037 | 0.99086 | 0.99096 | 0.99118
|  |  |  | 155 | 0.80557 | 0.80557 | 0.99499 | 0.98373 | 0.98786 | 0.98677 | 0.9863 | 0.98617 | 0.98651 | 0.98595 | 0.98649 | 0.98675 | 0.98734 | 0.9872
|  |  |  | 160 | 0.73496 | 0.73496 | 0.96768 | 0.97998 | 0.97881 | 0.98031 | 0.98051 | 0.98054 | 0.98082 | 0.98101 | 0.98147 | 0.98156 | 0.98184 | 0.98241
|  |  |  | 165 | 0.93938 | 0.93938 | 0.95276 | 0.96614 | 0.97026 | 0.96946 | 0.9695 | 0.96927 | 0.96962 | 0.97025 | 0.97034 | 0.97106 | 0.97203 | 0.97273
|  |  |  | 170 | 0.92455 | 0.92455 | 0.9398 | 0.95504 | 0.95922 | 0.95367 | 0.95434 | 0.9554 | 0.95643 | 0.9573 | 0.958 | 0.95951 | 0.96052 | 0.9622
|  |  |  | 172 | 0.92347 | 0.92347 | 0.93379 | 0.94411 | 0.95443 | 0.94399 | 0.94446 | 0.94551 | 0.94712 | 0.94868 | 0.94986 | 0.95088 | 0.95195 | 0.95375
|  |  |  | 174 | 0.91726 | 0.91726 | 0.92647 | 0.93569 | 0.94491 | 0.93408 | 0.93428 | 0.93573 | 0.93719 | 0.93876 | 0.94067 | 0.94247 | 0.94373 | 0.94652
|  |  |  | 175 | 0.97072 | 0.97072 | 0.95975 | 0.94877 | 0.9378 | 0.92683 | 0.92817 | 0.93013 | 0.93188 | 0.93356 | 0.93561 | 0.93774 | 0.9394 | 0.94243
|  |  |  | 176 | 0.97252 | 0.97252 | 0.95978 | 0.94704 | 0.9343 | 0.92156 | 0.9232 | 0.92479 | 0.92762 | 0.92942 | 0.9315 | 0.93393 | 0.9361 | 0.93885
|  |  |  | 177 | 0.96567 | 0.96567 | 0.95405 | 0.94243 | 0.9308 | 0.91918 | 0.91904 | 0.92083 | 0.92391 | 0.92591 | 0.92798 | 0.92973 | 0.93223 | 0.93609
|  |  |  | 178 | 0.99056 | 0.99056 | 0.97109 | 0.95163 | 0.93216 | 0.91269 | 0.91387 | 0.91782 | 0.91995 | 0.92215 | 0.9243 | 0.92546 | 0.92864 | 0.93277
|  |  |  | 179 | 1.0055 | 1.0055 | 0.98109 | 0.95669 | 0.9323 | 0.90791 | 0.90822 | 0.91092 | 0.91257 | 0.91546 | 0.91813 | 0.91837 | 0.92066 | 0.92586
|  |  |  | 180 | 0.99645 | 0.99645 | 0.97454 | 0.95263 | 0.93072 | 0.90882 | 0.90663 | 0.90661 | 0.90857 | 0.91175 | 0.91456 | 0.91365 | 0.91442 | 0.92149