## General Information
| General Information | 125I - Best Medical model 2301
Source Information
Source name: | Best Medical model 2301
Radionuclide: | I-125
Model Number: | 2301
Distributed by: | BestMedical International
Manufactured by:
References
1) | "Update of AAPM Task Group No. 43 Report: A revised AAPM protocol for brachytherapy dose calculations"
Med Phys 31 (3), March 2004, 633-674
M.J. Rivard, B.M. Coursey, L.A. DeWerd, W.F. Hanson, M.S. Huq, G.S. Ibbott, M.G. Mitch, R. Nath,J.F. Williamson
2) | "Erratum: "Update of AAPM Task Group No. 43 Report: AAPM protocol for brachytherapy dose calculations" [Med. Phys. 31, 633-674 (2004)]""
Med Phys 31(12), December 2004, 3532-3533
M.J. Rivard, B.M. Coursey, L.A. DeWerd, W.F. Hanson, M.S. Huq, G.S. Ibbott, M.G. Mitch, R. Nath,J.F. Williamson
Calculation details
Method: | Monte Carlo and TLD averaged
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised
2006-01-26

## Source design and materials
Source design and materials | 125I - Best Medical model 2301

## Geometry Function
Geometry Function |  |  |  | 125I - Best Medical model 2301
Active source length,
| L = | 4 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | 125I - Best Medical model 2301
L = | 1.018 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | 125I - Best Medical model 2301
F(r,q)
r (cm)
q (º) | 1 | 2 | 3 | 4 | 5 | 6 | 7
0 | 0.867 | 0.854 | 0.922 | 0.902 | 0.894 | 0.893 | 0.858
5 | 0.724 | 0.72 | 0.726 | 0.738 | 0.753 | 0.771 | 0.8
10 | 0.653 | 0.671 | 0.699 | 0.727 | 0.732 | 0.764 | 0.782
20 | 0.785 | 0.794 | 0.809 | 0.814 | 0.825 | 0.852 | 0.821
30 | 0.9 | 0.89 | 0.885 | 0.892 | 0.899 | 0.915 | 0.873
40 | 0.982 | 0.954 | 0.947 | 0.939 | 0.943 | 0.976 | 0.937
50 | 1.014 | 0.992 | 0.985 | 0.991 | 0.997 | 0.989 | 0.961
60 | 1.03 | 1.01 | 1.009 | 1.007 | 1.01 | 1.019 | 1.002
70 | 1.036 | 1.026 | 1.016 | 1.023 | 1.011 | 1.035 | 1.01
80 | 1.04 | 1.03 | 1.019 | 1.017 | 1.01 | 1.02 | 1.005

## Radial Dose Function
Radial-Dose Function Table |  |  | 125I - Best Medical model 2301
Line source aproximation,
L = | 4 | mm | Point source aproximation,
r (cm) | g(r) |  | r (cm) | g(r)
0.1 | 1.033 |  | 0.1 | 0.579
0.15 | 1.029 |  | 0.15 | 0.725
0.25 | 1.027 |  | 0.25 | 0.878
0.5 | 1.028 |  | 0.5 | 0.991
0.75 | 1.03 |  | 0.75 | 1.02
1 | 1 |  | 1 | 1
1.5 | 0.938 |  | 1.5 | 0.945
2 | 0.866 |  | 2 | 0.875
3 | 0.707 |  | 3 | 0.715
4 | 0.555 |  | 4 | 0.562
5 | 0.427 |  | 5 | 0.432
6 | 0.32 |  | 6 | 0.324
7 | 0.248 |  | 7 | 0.251
8 | 0.187 |  | 8 | 0.189
9 | 0.142 |  | 9 | 0.144
10 | 0.103 |  | 10 | 0.104

## Anisotropy Factor
Anisotropy Factor |  |  | 125I - Best Medical model 2301
| r (cm) | fan(r)
| 1 | 0.986
| 2 | 0.976
| 3 | 0.968
| 4 | 0.971
| 5 | 0.969
| 6 | 0.991
| 7 | 0.969

## QA
Dose rates for QA |  |  | 125I - Best Medical model 2301
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources using
the 1-D formalism of Eq. (11) with interpolation for gL(r) and Øan(r).
| r (cm) | fan(r)
| 0.5 | 3.813
| 1 | 0.962
| 1.5 | 0.413
| 2 | 0.22
| 3 | 0.0783
| 4 | 0.0347
| 5 | 0.0171
| 6 | 0.00908
| 7 | 0.00506