## General Information
| General Information | I-125 - International Brachytherapy model 1251L source
Source Information
Source name: | International Brachytherapy Intersource
Radionuclide: | I-125
Model Number: | 1251L source
Distributed by:
Manufactured by:
References
1) | "Supplement to the 2004 update of the AAPM Task Group Nº43 Report"
Med Phys 34(6), June 2007(2187-2205)
M.J. Rivard, W.M.Butler, L.A. DeWerd, M.S. Huq, G.S Ibbott, A.S. Meigooni, C.S. Melhus, M G. Mitch, R. Nath, J. F. Williamson
Calculation details
Method:
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised

## Geometry and materials
Geometry and materials | I-125 - International Brachytherapy model 1251L source

## Geometry Function
Geometry Function |  |  | I-125 - International Brachytherapy model 1251L source
Source length for Line source approximation:
L = | 4.35 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | I-125 - International Brachytherapy model 1251L source
L = | 1.038 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | I-125 - International Brachytherapy model 1251L source
F(r,q)
r (cm)
q (º) | 0.5 | 1 | 2 | 3 | 5
0 | 0.476 | 0.544 | 0.653 | 0.68 | 0.703
5 | 0.645 | 0.626 | 0.656 | 0.713 | 0.718
10 | 0.725 | 0.699 | 0.709 | 0.736 | 0.751
20 | 0.81 | 0.783 | 0.789 | 0.81 | 0.817
30 | 0.867 | 0.849 | 0.849 | 0.859 | 0.854
40 | 0.923 | 0.9 | 0.91 | 0.911 | 0.911
50 | 0.966 | 0.946 | 0.946 | 0.949 | 0.954
60 | 0.991 | 0.979 | 0.971 | 0.976 | 0.968
70 | 0.998 | 0.988 | 0.991 | 0.996 | 0.988
80 | 1.002 | 0.996 | 0.997 | 0.995 | 0.988
90 | 1 | 1 | 1 | 1 | 1

## Radial-Dose Function Table
Radial-Dose Function Table |  |  | I-125 - International Brachytherapy model 1251L source
Line source approximation,
L = | 4.35 | mm | Point source approximation
r (cm) | gL(r) |  | r (cm) | gp(r)
0.1 | 0.757 |  | 0.1 | 0.403
0.15 | 0.841 |  | 0.15 | 0.569
0.25 | 0.963 |  | 0.25 | 0.805
0.5 | 1.021 |  | 0.5 | 0.978
0.75 | 1.024 |  | 0.75 | 1.012
1 | 1 |  | 1 | 1
1.5 | 0.937 |  | 1.5 | 0.945
2 | 0.859 |  | 2 | 0.869
3 | 0.7 |  | 3 | 0.71
4 | 0.554 |  | 4 | 0.562
5 | 0.425 |  | 5 | 0.432
6 | 0.323 |  | 6 | 0.328
7 | 0.24 |  | 7 | 0.244
8 | 0.18 |  | 8 | 0.183
9 | 0.138 |  | 9 | 0.141
10 | 0.101 |  | 10 | 0.102

## Anisotropy Factor
Anisotropy Factor |  |  | I-125 - International Brachytherapy model 1251L source
| r (cm) | fan(r)
| 0.1 | 1.162
| 0.15 | 1.327
| 0.25 | 1.296
| 0.5 | 1.028
| 0.75 | 0.992
| 1 | 0.958
| 2 | 0.945
| 3 | 0.948
| 5 | 0.945

## QA
Dose rates for QA |  |  | I-125 - International Brachytherapy model 1251L source
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources included in this report using gL(r) and Øan(r),
and the 1-D formalism of Eq. (11) from the 2004 AAPM TG-43U1. Results using interpolated gL(r) or Øan(r) data are highlighted in boldface while
extrapolated results based on gL(r) and/or Øan(r) data are underlined.
| r (cm) | fan(r)
| 0.1 | 48.6
| 0.15 | 34.9
| 0.25 | 17.3
| 0.5 | 4.18
| 0.75 | 1.85
| 1 | 0.994
| 1.5 | 0.415
| 2 | 0.213
| 3 | 0.0776
| 4 | 0.0345
| 5 | 0.0169
| 6 | 0.00893
| 7 | 0.00488
| 8 | 0.00281
| 9 | 0.0017
| 10 | 0.001