## General Information
| General Information | I-125 - Mills Biopharmaceuticals Corporation model SL-125/SH-125
Source Information
Source name: | Mills Biopharmaceuticals Corporation model SL-125/SH-125
Radionuclide: | I-125
Model Number: | model SL-125/SH-125
Distributed by:
Manufactured by:
References
1) | "Supplement to the 2004 update of the AAPM Task Group Nº43 Report"
Med Phys 34(6), June 2007(2187-2205)
M.J. Rivard, W.M.Butler, L.A. DeWerd, M.S. Huq, G.S Ibbott, A.S. Meigooni, C.S. Melhus, M G. Mitch, R. Nath, J. F. Williamson
Calculation details
Method:
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised

## Geometry and materials
Geometry and materials | I-125 - Mills Biopharmaceuticals Corporation model SL-125/SH-125

## Geometry Function
Geometry Function |  |  | I-125 - Mills Biopharmaceuticals Corporation model SL-125/SH-125
Source length for Line source approximation:
L = | 3 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | I-125 - Mills Biopharmaceuticals Corporation model SL-125/SH-125
L = | 0.953 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | I-125 - Mills Biopharmaceuticals Corporation model SL-125/SH-125
F(r,q)
r (cm)
q (º) | 1 | 2 | 3 | 4 | 5
0 | 0.359 | 0.424 | 0.471 | 0.501 | 0.52
10 | 0.429 | 0.493 | 0.535 | 0.563 | 0.574
20 | 0.568 | 0.61 | 0.643 | 0.672 | 0.67
30 | 0.71 | 0.744 | 0.759 | 0.771 | 0.762
40 | 0.823 | 0.842 | 0.852 | 0.863 | 0.857
50 | 0.918 | 0.926 | 0.936 | 0.937 | 0.921
60 | 0.973 | 0.972 | 0.98 | 0.986 | 0.974
70 | 0.985 | 0.987 | 0.989 | 0.993 | 0.993
80 | 0.991 | 1 | 1.013 | 1.002 | 0.993
90 | 1 | 1 | 1 | 1 | 1

## Radial-Dose Function Table
Radial-Dose Function Table |  |  | I-125 - Mills Biopharmaceuticals Corporation model SL-125/SH-125
Line source approximation,
L = | 3 | mm | Point source approximation
r (cm) | gL(r) |  | r (cm) | gP(r)
0.1 | 1.101 |  | 0.1 | 0.727
0.15 | 1.101 |  | 0.15 | 0.871
0.25 | 1.101 |  | 0.25 | 0.999
0.5 | 1.084 |  | 0.5 | 1.061
0.75 | 1.041 |  | 0.75 | 1.035
1 | 1 |  | 1 | 1
1.5 | 0.898 |  | 1.5 | 0.901
2 | 0.795 |  | 2 | 0.799
3 | 0.61 |  | 3 | 0.614
4 | 0.456 |  | 4 | 0.459
5 | 0.338 |  | 5 | 0.34
6 | 0.25 |  | 6 | 0.252
7 | 0.183 |  | 7 | 0.184
8 | 0.134 |  | 8 | 0.135
9 | 0.098 |  | 9 | 0.099
10 | 0.072 |  | 10 | 0.072

## Anisotropy Factor
Anisotropy Factor |  |  | I-125 - Mills Biopharmaceuticals Corporation model SL-125/SH-125
| r (cm) | fan(r)
| 0.1 | 1.091
| 0.15 | 1.159
| 0.25 | 1.035
| 0.5 | 0.927
| 0.75 | 0.907
| 1 | 0.9
| 2 | 0.907
| 3 | 0.916
| 4 | 0.921
| 5 | 0.914

## QA
Dose rates for QA |  |  | I-125 - Mills Biopharmaceuticals Corporation model SL-125/SH-125
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources included in this report using gL(r) and Øan(r),
and the 1-D formalism of Eq. (11) from the 2004 AAPM TG-43U1. Results using interpolated gL(r) or Øan(r) data are highlighted in boldface while
extrapolated results based on gL(r) and/or Øan(r) data are underlined.
| r (cm) | fan(r)
| 0.1 | 75.6
| 0.15 | 42.8
| 0.25 | 15.8
| 0.5 | 3.75
| 0.75 | 1.59
| 1 | 0.858
| 1.5 | 0.345
| 2 | 0.175
| 3 | 0.0596
| 4 | 0.0252
| 5 | 0.0119
| 6 | 0.00609
| 7 | 0.00328
| 8 | 0.00184
| 9 | 0.00106
| 10 | 0.00063