## General Information
| General Information | I-125 - Source Tech Medical STM1251
Source Information
Source name: | Source Tech Medical STM1251
Radionuclide: | 125-I
Model Number: | STM 1252 source
Distributed by:
Manufactured by:
References
1) | "Supplement to the 2004 update of the AAPM Task Group Nº43 Report"
Med Phys 34(6), June 2007(2187-2205)
M.J. Rivard, W.M.Butler, L.A. DeWerd, M.S. Huq, G.S Ibbott, A.S. Meigooni, C.S. Melhus, M G. Mitch, R. Nath, J. F. Williamson
Calculation details
Method:
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised

## Geometry and materials
Geometry and materials | I-125 - Source Tech Medical STM1251

## Geometry Function
Geometry Function |  |  | I-125 - Source Tech Medical STM1251
Source length for Line source approximation:
L = | 3.81 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | I-125 - Source Tech Medical STM1251
L = | 1.018 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | I-125 - Source Tech Medical STM1251
F(r,q)
r (cm)
q (º) | 0.25 | 0.5 | 1 | 2 | 3 | 5 | 7
0 | 0.863 | 0.524 | 0.423 | 0.453 | 0.5 | 0.564 | 0.607
2 | 0.865 | 0.489 | 0.616 | 0.701 | 0.702 | 0.706 | 0.72
5 | 0.784 | 0.668 | 0.599 | 0.611 | 0.637 | 0.657 | 0.682
7 | 0.861 | 0.588 | 0.575 | 0.603 | 0.632 | 0.655 | 0.682
10 | 0.778 | 0.562 | 0.579 | 0.617 | 0.649 | 0.672 | 0.7
20 | 0.889 | 0.688 | 0.698 | 0.722 | 0.75 | 0.761 | 0.781
30 | 0.949 | 0.816 | 0.808 | 0.819 | 0.841 | 0.838 | 0.845
40 | 0.979 | 0.898 | 0.888 | 0.891 | 0.903 | 0.901 | 0.912
50 | 0.959 | 0.956 | 0.943 | 0.941 | 0.95 | 0.941 | 0.945
60 | 0.98 | 0.988 | 0.982 | 0.98 | 0.985 | 0.973 | 0.982
70 | 0.989 | 0.973 | 1.005 | 1.002 | 1.011 | 0.995 | 0.998
80 | 0.994 | 0.994 | 0.989 | 1.015 | 1.018 | 1.003 | 1.011
90 | 1 | 1 | 1 | 1 | 1 | 1 | 1

## Radial-Dose Function Table
Radial-Dose Function Table |  |  | I-125 - Source Tech Medical STM1251
Line source approximation,
L = | 3.81 | mm | Point source approximation
r (cm) | gL(r) |  | r (cm) | gP(r)
0.1 | 0.941 |  | 0.1 | 0.544
0.15 | 0.972 |  | 0.15 | 0.7
0.25 | 1.013 |  | 0.25 | 0.876
0.5 | 1.033 |  | 0.5 | 0.999
0.75 | 1.022 |  | 0.75 | 1.013
1 | 1 |  | 1 | 1
1.5 | 0.937 |  | 1.5 | 0.943
2 | 0.856 |  | 2 | 0.864
3 | 0.691 |  | 3 | 0.698
4 | 0.54 |  | 4 | 0.546
5 | 0.415 |  | 5 | 0.42
6 | 0.314 |  | 6 | 0.318
7 | 0.236 |  | 7 | 0.239
8 | 0.176 |  | 8 | 0.178
9 | 0.131 |  | 9 | 0.133
10 | 0.0969 |  | 10 | 0.098

## Anisotropy Factor
Anisotropy Factor |  |  | I-125 - Source Tech Medical STM1251
| r (cm) | fan(r)
| 0.1 | 1.172
| 0.15 | 1.317
| 0.25 | 1.21
| 0.5 | 0.982
| 0.75 | 0.962
| 1 | 0.942
| 2 | 0.937
| 3 | 0.947
| 5 | 0.938
| 7 | 0.944

## QA
Dose rates for QA |  |  | I-125 - Source Tech Medical STM1251
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance for the 8 brachytherapy sources included in this report using gL(r) and Øan(r),
and the 1-D formalism of Eq. (11) from the 2004 AAPM TG-43U1. Results using interpolated gL(r) or Øan(r) data are highlighted in boldface while
extrapolated results based on gL(r) and/or Øan(r) data are underlined.
| r (cm) | fan(r)
| 0.1 | 64.8
| 0.15 | 41.6
| 0.25 | 17.3
| 0.5 | 3.99
| 0.75 | 1.76
| 1 | 0.958
| 1.5 | 0.401
| 2 | 0.206
| 3 | 0.0747
| 4 | 0.0327
| 5 | 0.016
| 6 | 0.00844
| 7 | 0.00468
| 8 | 0.00267
| 9 | 0.00157
| 10 | 0.000941