## Ir2.A85-2
| TG-43 CONSENSUS | [5 empty] | E&Z BEBIG 192Ir Ir2.A85-2
Dose calculation for photon-emitting brachytherapy sources with average energy higher than 50 keV: Full Report of the AAPM and ESTRO
Med. Phys. 39 (2012) 2904-2929
| L = | 1.109 |  | cGy/(h U)
Interpolated / extrapolated data are boldface / underlined. Values inside the source are in italics.
| gL(r) |  |  |  | F(r,q) | [13 empty] | QA Along-away
| L = | 0.35 | cm |  |  |  |  | Distance from Active Source Center (cm)
| r (cm) | gL( r) |  | Theta/deg | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 10 | z / cm | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7
| 0 | 0.99 |  | 0 | 0.666 | 0.666 | 0.634 | 0.602 | 0.592 | 0.59 | 0.596 | 0.617 | 0.642 | 0.664 | 0.683 | 0.701 | 0.72 | 0.753 | 7 | 0.015731 | 0.016439 | 0.016833 | 0.017153 | 0.017403 | 0.017739 | 0.017765 | 0.01699 | 0.015539 | 0.01376 | 0.012015 | 0.010381
| 0.25 | 0.99 |  | 1 | 0.667 | 0.667 | 0.637 | 0.607 | 0.597 | 0.599 | 0.608 | 0.632 | 0.656 | 0.678 | 0.697 | 0.714 | 0.732 | 0.761 | 6 | 0.02108 | 0.022181 | 0.022822 | 0.023371 | 0.023851 | 0.024282 | 0.024091 | 0.022435 | 0.019836 | 0.017097 | 0.014471 | 0.012201
| 0.5 | 0.996 |  | 2 | 0.667 | 0.667 | 0.641 | 0.615 | 0.609 | 0.616 | 0.629 | 0.654 | 0.678 | 0.699 | 0.716 | 0.733 | 0.749 | 0.775 | 5 | 0.02976 | 0.031677 | 0.032795 | 0.033772 | 0.034489 | 0.034906 | 0.034124 | 0.030435 | 0.025724 | 0.02124 | 0.017414 | 0.014244
| 0.75 | 0.998 |  | 3 | 0.665 | 0.665 | 0.644 | 0.623 | 0.623 | 0.631 | 0.643 | 0.666 | 0.689 | 0.71 | 0.726 | 0.743 | 0.758 | 0.784 | 4 | 0.045171 | 0.048702 | 0.051118 | 0.053084 | 0.05408 | 0.053721 | 0.050948 | 0.042387 | 0.033622 | 0.026365 | 0.02072 | 0.016468
| 1 | 1 |  | 4 | 0.665 | 0.665 | 0.65 | 0.635 | 0.635 | 0.644 | 0.655 | 0.679 | 0.699 | 0.72 | 0.736 | 0.753 | 0.766 | 0.791 | 3 | 0.077446 | 0.085803 | 0.091827 | 0.09516 | 0.095639 | 0.090432 | 0.080786 | 0.060173 | 0.043821 | 0.032261 | 0.024246 | 0.018661
| 1.5 | 1.002 |  | 5 | 0.668 | 0.668 | 0.658 | 0.648 | 0.648 | 0.658 | 0.669 | 0.691 | 0.712 | 0.731 | 0.747 | 0.763 | 0.776 | 0.798 | 2 | 0.16887 | 0.19514 | 0.21143 | 0.21249 | 0.20284 | 0.16977 | 0.13569 | 0.084952 | 0.055493 | 0.03822 | 0.027538 | 0.020589
| 2 | 1.004 |  | 6 | 0.679 | 0.679 | 0.67 | 0.661 | 0.663 | 0.673 | 0.683 | 0.705 | 0.724 | 0.742 | 0.757 | 0.772 | 0.785 | 0.805 | 1.5 | 0.2984 | 0.35651 | 0.37844 | 0.3603 | 0.32239 | 0.24095 | 0.1761 | 0.098895 | 0.061148 | 0.040835 | 0.028888 | 0.021356
| 3 | 1.005 |  | 8 | 0.705 | 0.705 | 0.698 | 0.691 | 0.692 | 0.702 | 0.712 | 0.731 | 0.748 | 0.764 | 0.778 | 0.79 | 0.801 | 0.82 | 1 | 0.68413 | 0.84688 | 0.81529 | 0.67985 | 0.54184 | 0.3394 | 0.22237 | 0.11191 | 0.065858 | 0.042929 | 0.029943 | 0.021936
| 4 | 1.003 |  | 10 | 0.732 | 0.732 | 0.728 | 0.724 | 0.725 | 0.733 | 0.742 | 0.759 | 0.774 | 0.787 | 0.8 | 0.81 | 0.819 | 0.836 | 0.5 | 3.2245 | 3.4457 | 2.1912 | 1.3542 | 0.88417 | 0.44564 | 0.26337 | 0.12147 | 0.069046 | 0.04427 | 0.030595 | 0.022295
| 5 | 0.999 |  | 15 | 0.806 | 0.806 | 0.8 | 0.794 | 0.795 | 0.8 | 0.805 | 0.816 | 0.826 | 0.836 | 0.845 | 0.851 | 0.859 | 0.87 | 0 | 379857899.2146 | 15.4814 | 4.2931 | 1.9526 | 1.109 | 0.49663 | 0.28046 | 0.12495 | 0.07018 | 0.044746 | 0.030829 | 0.022423
| 6 | 0.991 |  | 20 | 0.86 | 0.86 | 0.853 | 0.846 | 0.845 | 0.848 | 0.852 | 0.86 | 0.867 | 0.874 | 0.879 | 0.885 | 0.89 | 0.898 | -0.5 | 2.9906 | 3.4373 | 2.1899 | 1.3544 | 0.88499 | 0.44563 | 0.26322 | 0.12136 | 0.069046 | 0.04427 | 0.030595 | 0.022286
| 8 | 0.968 |  | 25 | 0.895 | 0.895 | 0.89 | 0.885 | 0.884 | 0.886 | 0.888 | 0.893 | 0.898 | 0.903 | 0.907 | 0.91 | 0.914 | 0.92 | -1 | 0.67604 | 0.84918 | 0.81493 | 0.67924 | 0.54151 | 0.3394 | 0.22239 | 0.11197 | 0.065858 | 0.042929 | 0.029941 | 0.02192
| 10 | 0.935 |  | 30 | 0.921 | 0.921 | 0.917 | 0.913 | 0.912 | 0.912 | 0.915 | 0.918 | 0.922 | 0.925 | 0.928 | 0.93 | 0.933 | 0.936 | -1.5 | 0.29435 | 0.35776 | 0.37892 | 0.36008 | 0.32221 | 0.24081 | 0.17598 | 0.098918 | 0.061148 | 0.040835 | 0.028885 | 0.021341
|  |  |  | 40 | 0.956 | 0.956 | 0.953 | 0.95 | 0.949 | 0.95 | 0.95 | 0.953 | 0.954 | 0.956 | 0.957 | 0.958 | 0.96 | 0.961 | -2 | 0.16774 | 0.19589 | 0.21196 | 0.21271 | 0.20277 | 0.16965 | 0.13556 | 0.08494 | 0.055493 | 0.03822 | 0.027536 | 0.020579
|  |  |  | 50 | 0.976 | 0.976 | 0.974 | 0.972 | 0.972 | 0.972 | 0.973 | 0.974 | 0.974 | 0.975 | 0.976 | 0.976 | 0.977 | 0.978 | -3 | 0.076943 | 0.085714 | 0.092046 | 0.095373 | 0.095787 | 0.090401 | 0.080719 | 0.060165 | 0.043821 | 0.032258 | 0.024258 | 0.018655
|  |  |  | 60 | 0.99 | 0.99 | 0.986 | 0.986 | 0.985 | 0.985 | 0.986 | 0.986 | 0.987 | 0.987 | 0.988 | 0.987 | 0.988 | 0.988 | -4 | 0.044748 | 0.048502 | 0.051294 | 0.053159 | 0.054193 | 0.053761 | 0.050939 | 0.042357 | 0.033604 | 0.026353 | 0.020729 | 0.016469
|  |  |  | 70 | 0.996 | 0.996 | 0.994 | 0.994 | 0.993 | 0.994 | 0.994 | 0.994 | 0.995 | 0.995 | 0.995 | 0.995 | 0.996 | 0.995 | -5 | 0.02985 | 0.031487 | 0.032845 | 0.033848 | 0.034529 | 0.034939 | 0.034143 | 0.030435 | 0.025724 | 0.021239 | 0.0174 | 0.014238
|  |  |  | 80 | 0.999 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | -6 | 0.021172 | 0.022112 | 0.022814 | 0.023429 | 0.023866 | 0.02431 | 0.024128 | 0.022446 | 0.019834 | 0.017083 | 0.014458 | 0.012189
|  |  |  | 90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | -7 | 0.015799 | 0.016374 | 0.01679 | 0.017155 | 0.017425 | 0.017757 | 0.017795 | 0.016998 | 0.015539 | 0.013754 | 0.012009 | 0.010375
|  |  |  | 100 | 1 | 1 | 0.999 | 0.999 | 0.998 | 0.998 | 0.998 | 0.998 | 0.999 | 0.999 | 0.999 | 0.998 | 0.999 | 0.999
|  |  |  | 110 | 0.996 | 0.996 | 0.994 | 0.994 | 0.994 | 0.994 | 0.994 | 0.995 | 0.995 | 0.995 | 0.995 | 0.995 | 0.995 | 0.995
|  |  |  | 120 | 0.991 | 0.991 | 0.986 | 0.986 | 0.986 | 0.986 | 0.986 | 0.986 | 0.987 | 0.987 | 0.988 | 0.988 | 0.988 | 0.989
|  |  |  | 130 | 0.974 | 0.974 | 0.973 | 0.972 | 0.971 | 0.972 | 0.972 | 0.973 | 0.974 | 0.975 | 0.975 | 0.976 | 0.976 | 0.977
|  |  |  | 140 | 0.955 | 0.955 | 0.952 | 0.949 | 0.948 | 0.949 | 0.95 | 0.952 | 0.954 | 0.955 | 0.957 | 0.958 | 0.959 | 0.961
|  |  |  | 150 | 0.922 | 0.922 | 0.917 | 0.912 | 0.911 | 0.912 | 0.914 | 0.917 | 0.921 | 0.925 | 0.928 | 0.93 | 0.933 | 0.936
|  |  |  | 155 | 0.887 | 0.887 | 0.886 | 0.885 | 0.884 | 0.885 | 0.888 | 0.893 | 0.898 | 0.903 | 0.907 | 0.911 | 0.914 | 0.92
|  |  |  | 160 | 0.849 | 0.849 | 0.848 | 0.847 | 0.846 | 0.849 | 0.853 | 0.861 | 0.868 | 0.874 | 0.881 | 0.885 | 0.891 | 0.898
|  |  |  | 165 | 0.791 | 0.791 | 0.793 | 0.795 | 0.797 | 0.801 | 0.807 | 0.818 | 0.828 | 0.837 | 0.846 | 0.853 | 0.86 | 0.871
|  |  |  | 170 | 0.7175 | 0.7175 | 0.721 | 0.7245 | 0.728 | 0.735 | 0.744 | 0.76 | 0.775 | 0.788 | 0.8 | 0.81 | 0.82 | 0.836
|  |  |  | 172 | 0.686 | 0.686 | 0.69 | 0.694 | 0.698 | 0.706 | 0.716 | 0.735 | 0.751 | 0.766 | 0.78 | 0.791 | 0.803 | 0.821
|  |  |  | 174 | 0.6505 | 0.6505 | 0.655 | 0.6595 | 0.664 | 0.673 | 0.684 | 0.706 | 0.726 | 0.744 | 0.759 | 0.772 | 0.785 | 0.806
|  |  |  | 175 | 0.6335 | 0.6335 | 0.638 | 0.6425 | 0.647 | 0.656 | 0.668 | 0.691 | 0.711 | 0.73 | 0.747 | 0.762 | 0.775 | 0.798
|  |  |  | 176 | 0.6175 | 0.6175 | 0.622 | 0.6265 | 0.631 | 0.64 | 0.652 | 0.676 | 0.697 | 0.717 | 0.735 | 0.751 | 0.765 | 0.79
|  |  |  | 177 | 0.5985 | 0.5985 | 0.604 | 0.6095 | 0.615 | 0.626 | 0.638 | 0.662 | 0.685 | 0.706 | 0.725 | 0.742 | 0.756 | 0.782
|  |  |  | 178 | 0.5915 | 0.5915 | 0.595 | 0.5985 | 0.602 | 0.609 | 0.622 | 0.647 | 0.67 | 0.693 | 0.713 | 0.73 | 0.746 | 0.771
|  |  |  | 179 | 0.5885 | 0.5885 | 0.589 | 0.5895 | 0.59 | 0.591 | 0.602 | 0.626 | 0.649 | 0.676 | 0.696 | 0.713 | 0.728 | 0.759
|  |  |  | 180 | 0.5895 | 0.5895 | 0.588 | 0.5865 | 0.585 | 0.582 | 0.592 | 0.613 | 0.636 | 0.666 | 0.686 | 0.704 | 0.718 | 0.755