## readme first
mHDR-v2 - mHDR-v2r - mHDR-v2c dosimetry datasets
Since the release of the version 4.5 of the Oncentra Brachy TPS, there have been three different TG-43 dosimetry datasets to choose from for the model mHDR-v2 Ir-192 brachytherapy source. The v2r dataset is from the publication by Granero et al. (2011), which updated the dose analysis by Daskalov et al (1998) for the same source (v2). Subsequently, the AAPM and ESTRO have issued consensus guidelines and TG-43 dosimetry datasets (Perez-Calatayud et al, 2012). In order to be compliant with these guidelines, the consensus TG-43 dosimetry dataset (v2c) was added in Oncentra Brachy 4.5 and is recommended by Elekta for clinical practice. The two other models are kept for historical reasons and for transitional period.

## Source geometry
Warning: the modelled sources in Daskalov et al. and Granero et al. (see references in HEBD report) are considered to be the same source model: mHDR-v2. So, only consensus data for the mHDR-v2 source model are provided.
Conclusions in Granero et al.

## mHDR-v2
| TG-43 CONSENSUS | [5 empty] | Nucletron HDR 192Ir mHDR-v2
Dose calculation for photon-emitting brachytherapy sources with average energy higher than 50 keV: Full Report of the AAPM and ESTRO
Med. Phys. 39 (2012) 2904-2929
| L = | 1.1094 |  | cGy/(h U)
Interpolated / extrapolated data are boldface / underlined. Values inside the source are in italics.
| gL(r) |  |  |  | F(r,q) | [19 empty] | QA Along-away
| L = | 0.35 | cm |  |  |  |  | Distance from Active Source Center (cm)
| r (cm) | gL( r) |  | Theta/deg | 0 | 0.06 | 0.08 | 0.1 | 0.15 | 0.2 | 0.25 | 0.3 | 0.35 | 0.4 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 8 | 10 | z / cm | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7
| 0 | 1.2756 |  | 0 | 0.95131 | 0.95131 | 0.93406 | 0.91682 | 0.87371 | 0.8306 | 0.78749 | 0.74438 | 0.71369 | 0.69163 | 0.61933 | 0.61047 | 0.61373 | 0.62537 | 0.65045 | 0.68894 | 0.71074 | 0.73274 | 0.76757 | 0.79785 | 7 | 0.016902 | 0.017092 | 0.017244 | 0.017486 | 0.017721 | 0.017999 | 0.01797 | 0.017129 | 0.015639 | 0.013852 | 0.012086 | 0.010439
| 0.06 | 1.2756 |  | 2 | 0.94742 | 0.94742 | 0.93047 | 0.91352 | 0.87114 | 0.82876 | 0.78638 | 0.744 | 0.71364 | 0.6927 | 0.639 | 0.63429 | 0.64014 | 0.65139 | 0.67485 | 0.70435 | 0.7246 | 0.74368 | 0.77497 | 0.80072 | 6 | 0.022741 | 0.023107 | 0.023463 | 0.023904 | 0.024278 | 0.024637 | 0.024363 | 0.022582 | 0.019957 | 0.017189 | 0.014549 | 0.012263
| 0.08 | 1.199 |  | 4 | 0.9435 | 0.9435 | 0.92685 | 0.91019 | 0.86855 | 0.82691 | 0.78527 | 0.74363 | 0.71358 | 0.69378 | 0.65866 | 0.65811 | 0.66656 | 0.6774 | 0.69925 | 0.71976 | 0.73847 | 0.75462 | 0.78237 | 0.80358 | 5 | 0.032003 | 0.032817 | 0.033656 | 0.034464 | 0.035066 | 0.035331 | 0.03445 | 0.030636 | 0.025874 | 0.021333 | 0.0175 | 0.01432
| 0.1 | 1.1098 |  | 6 | 1.0588 | 1.0588 | 1.0333 | 1.0079 | 0.94435 | 0.88079 | 0.81723 | 0.75367 | 0.72115 | 0.7069 | 0.68406 | 0.68465 | 0.69252 | 0.70256 | 0.72214 | 0.74067 | 0.75747 | 0.77181 | 0.79598 | 0.81581 | 4 | 0.048706 | 0.050465 | 0.052437 | 0.05405 | 0.054932 | 0.054224 | 0.05128 | 0.042625 | 0.03379 | 0.026482 | 0.020817 | 0.016559
| 0.15 | 1.0183 |  | 8 | 0.99899 | 0.99899 | 0.98006 | 0.96112 | 0.91378 | 0.86644 | 0.8191 | 0.77176 | 0.74441 | 0.73 | 0.71049 | 0.7115 | 0.71936 | 0.7286 | 0.74628 | 0.76256 | 0.77684 | 0.7902 | 0.81242 | 0.83002 | 3 | 0.081943 | 0.088597 | 0.093648 | 0.096537 | 0.096689 | 0.090968 | 0.081263 | 0.060445 | 0.04401 | 0.032423 | 0.024361 | 0.018749
| 0.2 | 1.0006 |  | 10 | 1.0071 | 1.0071 | 0.98923 | 0.97135 | 0.92665 | 0.88195 | 0.83725 | 0.79255 | 0.76597 | 0.75485 | 0.7385 | 0.73934 | 0.74619 | 0.75434 | 0.76994 | 0.78452 | 0.7973 | 0.809 | 0.82888 | 0.84384 | 2 | 0.17762 | 0.2006 | 0.21434 | 0.21404 | 0.20369 | 0.17043 | 0.13602 | 0.085295 | 0.055749 | 0.038405 | 0.027677 | 0.020683
| 0.25 | 0.99471 |  | 12 | 1.0065 | 1.0065 | 0.99088 | 0.97524 | 0.93614 | 0.89704 | 0.85794 | 0.81884 | 0.79147 | 0.7797 | 0.76527 | 0.76614 | 0.7719 | 0.77887 | 0.79238 | 0.80515 | 0.81594 | 0.82631 | 0.84365 | 0.85677 | 1.5 | 0.31118 | 0.36444 | 0.38213 | 0.3616 | 0.3236 | 0.24135 | 0.17647 | 0.099302 | 0.061419 | 0.041033 | 0.029049 | 0.021457
| 0.5 | 0.9967 |  | 14 | 1.158 | 1.158 | 1.1288 | 1.0996 | 1.0267 | 0.95378 | 0.88085 | 0.83005 | 0.81067 | 0.80359 | 0.78962 | 0.79019 | 0.79535 | 0.80148 | 0.81324 | 0.82466 | 0.83485 | 0.84375 | 0.85852 | 0.87055 | 1 | 0.70744 | 0.85889 | 0.81768 | 0.68162 | 0.54227 | 0.33997 | 0.22291 | 0.11236 | 0.066168 | 0.043145 | 0.030122 | 0.022038
| 0.75 | 0.99833 |  | 16 | 1.2692 | 1.2692 | 1.2304 | 1.1915 | 1.0943 | 0.99717 | 0.9 | 0.85081 | 0.83167 | 0.82485 | 0.81185 | 0.81201 | 0.8163 | 0.82161 | 0.83198 | 0.84161 | 0.84998 | 0.85764 | 0.87063 | 0.8809 | 0.5 | 3.4461 | 3.4699 | 2.1928 | 1.3544 | 0.88518 | 0.44637 | 0.26387 | 0.12187 | 0.069379 | 0.044484 | 0.030772 | 0.02239
| 1 | 1 |  | 18 | 1.3784 | 1.3784 | 1.3296 | 1.2808 | 1.1588 | 1.0368 | 0.91473 | 0.86713 | 0.84987 | 0.84375 | 0.83143 | 0.8315 | 0.83546 | 0.84022 | 0.8491 | 0.85723 | 0.86496 | 0.87132 | 0.88243 | 0.89171 | 0 | 680353045.8707 | 15.467 | 4.291 | 1.9532 | 1.1094 | 0.49736 | 0.28097 | 0.12538 | 0.070506 | 0.04495 | 0.030998 | 0.022512
| 1.5 | 1.0028 |  | 20 | 1.7837 | 1.7837 | 1.6776 | 1.5715 | 1.3063 | 1.0411 | 0.93297 | 0.88543 | 0.86823 | 0.8606 | 0.84907 | 0.84891 | 0.85197 | 0.85609 | 0.86365 | 0.87107 | 0.87754 | 0.88341 | 0.89323 | 0.90111 | -0.5 | 2.621 | 3.4659 | 2.1962 | 1.355 | 0.88542 | 0.44653 | 0.2639 | 0.1219 | 0.069378 | 0.044484 | 0.03077 | 0.022388
| 2 | 1.005 |  | 22 | 1.7843 | 1.7843 | 1.6795 | 1.5746 | 1.3125 | 1.0504 | 0.94173 | 0.89346 | 0.88079 | 0.87528 | 0.86426 | 0.86425 | 0.86717 | 0.87071 | 0.87733 | 0.88375 | 0.88956 | 0.89442 | 0.90324 | 0.91027 | -1 | 0.60831 | 0.84858 | 0.81588 | 0.68136 | 0.54294 | 0.33995 | 0.22295 | 0.11236 | 0.066172 | 0.043146 | 0.030118 | 0.022034
| 3 | 1.0075 |  | 26 | 1.7038 | 1.7038 | 1.6099 | 1.516 | 1.2812 | 1.0464 | 0.95344 | 0.91961 | 0.90571 | 0.89967 | 0.89077 | 0.89039 | 0.89236 | 0.89507 | 0.89998 | 0.90463 | 0.90921 | 0.91286 | 0.91968 | 0.92485 | -1.5 | 0.27357 | 0.35487 | 0.37975 | 0.36077 | 0.32324 | 0.24159 | 0.17644 | 0.099312 | 0.061429 | 0.041037 | 0.029047 | 0.021453
| 4 | 1.0067 |  | 30 | 1.0889 | 1.0889 | 1.1192 | 1.1494 | 1.225 | 1.0491 | 0.96146 | 0.93248 | 0.92312 | 0.91929 | 0.91139 | 0.91097 | 0.91259 | 0.91457 | 0.91858 | 0.92236 | 0.92572 | 0.92868 | 0.93374 | 0.93763 | -2 | 0.15865 | 0.19274 | 0.2119 | 0.21299 | 0.20324 | 0.1703 | 0.13612 | 0.0853 | 0.055755 | 0.03841 | 0.027678 | 0.020682
| 5 | 1.0026 |  | 32 | 1.1573 | 1.1573 | 1.1674 | 1.1775 | 1.2028 | 1.0389 | 0.96631 | 0.9385 | 0.93083 | 0.92728 | 0.92059 | 0.92022 | 0.92139 | 0.92329 | 0.92669 | 0.92983 | 0.93267 | 0.93537 | 0.9399 | 0.94356 | -3 | 0.074516 | 0.083988 | 0.091525 | 0.095497 | 0.096166 | 0.090745 | 0.081167 | 0.060489 | 0.044021 | 0.032429 | 0.024364 | 0.018751
| 6 | 0.99552 |  | 36 | 1.1813 | 1.1813 | 1.1757 | 1.1701 | 1.156 | 1.0227 | 0.97115 | 0.94931 | 0.94408 | 0.94145 | 0.93546 | 0.93527 | 0.93596 | 0.93748 | 0.94024 | 0.94281 | 0.9448 | 0.94677 | 0.95 | 0.95281 | -4 | 0.042245 | 0.047152 | 0.050736 | 0.053115 | 0.054342 | 0.053984 | 0.051168 | 0.042595 | 0.033811 | 0.026486 | 0.020818 | 0.016559
| 8 | 0.9724 |  | 40 | 0.95373 | 0.95373 | 1.053 | 1.1522 | 1.1088 | 1.0157 | 0.97372 | 0.96138 | 0.95526 | 0.95283 | 0.948 | 0.948 | 0.94881 | 0.94995 | 0.95214 | 0.954 | 0.95596 | 0.9575 | 0.95994 | 0.96204 | -5 | 0.027949 | 0.030391 | 0.032359 | 0.033688 | 0.034577 | 0.035101 | 0.034327 | 0.030586 | 0.025853 | 0.021343 | 0.017498 | 0.014319
| 10 | 0.93894 |  | 50 | 1.0369 | 1.0369 | 1.0706 | 1.1042 | 1.0466 | 0.99887 | 0.98123 | 0.97579 | 0.97425 | 0.97293 | 0.97076 | 0.9703 | 0.97122 | 0.97203 | 0.97299 | 0.97374 | 0.97446 | 0.97518 | 0.97663 | 0.9778 | -6 | 0.020046 | 0.021343 | 0.022467 | 0.023257 | 0.023838 | 0.024389 | 0.024237 | 0.02254 | 0.019935 | 0.017169 | 0.014554 | 0.012263
|  |  |  | 60 | 1.0085 | 1.0085 | 1.0412 | 1.0619 | 1.0128 | 0.99771 | 0.99258 | 0.98738 | 0.98586 | 0.98532 | 0.98449 | 0.98477 | 0.98546 | 0.98599 | 0.98637 | 0.98684 | 0.9873 | 0.98761 | 0.98803 | 0.98855 | -7 | 0.014999 | 0.015761 | 0.016471 | 0.016974 | 0.017351 | 0.017798 | 0.017858 | 0.017077 | 0.01561 | 0.01384 | 0.012074 | 0.01044
|  |  |  | 70 | 1.0782 | 1.0782 | 1.0227 | 1.026 | 1.0014 | 0.99717 | 0.99553 | 0.99471 | 0.99424 | 0.99366 | 0.99328 | 0.99361 | 0.99393 | 0.99414 | 0.99462 | 0.99474 | 0.9949 | 0.99491 | 0.9952 | 0.99544
|  |  |  | 80 | 1.0198 | 1.0198 | 1.0052 | 1.0074 | 0.99958 | 1.0019 | 1.0028 | 0.99845 | 0.99846 | 0.99844 | 0.99814 | 0.99792 | 0.99853 | 0.99878 | 0.99887 | 0.99924 | 0.99945 | 0.99961 | 0.99976 | 0.99958
|  |  |  | 90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  |  | 100 | 1.0115 | 1.0115 | 1.0023 | 1.0081 | 0.99602 | 0.99544 | 0.99897 | 0.99865 | 0.9985 | 0.9986 | 0.99814 | 0.99819 | 0.99839 | 0.9987 | 0.99911 | 0.99922 | 0.99948 | 0.99949 | 0.99942 | 0.99942
|  |  |  | 110 | 1.0687 | 1.0687 | 1.0292 | 1.0247 | 1.0063 | 0.99442 | 0.99578 | 0.99494 | 0.99428 | 0.99386 | 0.99359 | 0.99389 | 0.99439 | 0.99453 | 0.99455 | 0.99492 | 0.99505 | 0.99495 | 0.99541 | 0.9953
|  |  |  | 120 | 1.0036 | 1.0036 | 1.0493 | 1.0602 | 1.0204 | 0.99929 | 0.992 | 0.98761 | 0.98643 | 0.98594 | 0.98504 | 0.98508 | 0.98544 | 0.98612 | 0.98654 | 0.9869 | 0.98743 | 0.98778 | 0.98802 | 0.98852
|  |  |  | 130 | 1.0556 | 1.0556 | 1.0801 | 1.1045 | 1.0467 | 1.003 | 0.98549 | 0.97632 | 0.97494 | 0.97387 | 0.97092 | 0.97084 | 0.97116 | 0.97169 | 0.97266 | 0.97389 | 0.97478 | 0.97536 | 0.97652 | 0.97768
|  |  |  | 132 | 1.0426 | 1.0426 | 1.0766 | 1.1105 | 1.0584 | 1.0024 | 0.98169 | 0.97376 | 0.97176 | 0.97039 | 0.96722 | 0.96696 | 0.96775 | 0.96819 | 0.96923 | 0.97043 | 0.97155 | 0.97231 | 0.97396 | 0.97463
|  |  |  | 134 | 1.0215 | 1.0215 | 1.0782 | 1.1348 | 1.0679 | 1.007 | 0.98196 | 0.97065 | 0.9684 | 0.96663 | 0.96316 | 0.96277 | 0.96348 | 0.96426 | 0.96549 | 0.96671 | 0.96805 | 0.96897 | 0.97048 | 0.97161
|  |  |  | 136 | 1.0108 | 1.0108 | 1.0746 | 1.1384 | 1.076 | 1.0138 | 0.97945 | 0.96684 | 0.96443 | 0.9626 | 0.9585 | 0.95845 | 0.95881 | 0.95966 | 0.96103 | 0.96239 | 0.96392 | 0.96474 | 0.9668 | 0.96847
|  |  |  | 138 | 1.0798 | 1.0798 | 1.1128 | 1.1457 | 1.098 | 1.0162 | 0.97729 | 0.9632 | 0.9605 | 0.95836 | 0.95355 | 0.95296 | 0.95388 | 0.95459 | 0.95643 | 0.95792 | 0.95966 | 0.9609 | 0.96288 | 0.96489
|  |  |  | 140 | 0.98342 | 0.98342 | 1.0683 | 1.1531 | 1.1134 | 1.0204 | 0.98201 | 0.96125 | 0.95573 | 0.95355 | 0.94787 | 0.94747 | 0.94831 | 0.94919 | 0.95123 | 0.95327 | 0.95498 | 0.95655 | 0.95883 | 0.96076
|  |  |  | 144 | 1.1836 | 1.1836 | 1.1763 | 1.169 | 1.1508 | 1.0328 | 0.97756 | 0.95098 | 0.94492 | 0.94221 | 0.93566 | 0.93502 | 0.93558 | 0.93688 | 0.93952 | 0.94205 | 0.94422 | 0.94638 | 0.94937 | 0.95199
|  |  |  | 148 | 1.1405 | 1.1405 | 1.1545 | 1.1686 | 1.2037 | 1.0313 | 0.97555 | 0.94233 | 0.93156 | 0.92779 | 0.91943 | 0.91929 | 0.92032 | 0.92164 | 0.92524 | 0.92852 | 0.93137 | 0.93426 | 0.93827 | 0.94199
|  |  |  | 150 | 1.0988 | 1.0988 | 1.1285 | 1.1581 | 1.2322 | 1.0517 | 0.96716 | 0.92968 | 0.92306 | 0.91953 | 0.91058 | 0.90995 | 0.91172 | 0.91285 | 0.91677 | 0.9205 | 0.92385 | 0.92678 | 0.93205 | 0.93601
|  |  |  | 154 | 1.631 | 1.631 | 1.5541 | 1.4772 | 1.2848 | 1.0925 | 0.95916 | 0.91371 | 0.90351 | 0.89923 | 0.88867 | 0.88834 | 0.89 | 0.89316 | 0.89754 | 0.90263 | 0.90718 | 0.91126 | 0.91771 | 0.92317
|  |  |  | 158 | 1.7254 | 1.7254 | 1.6362 | 1.547 | 1.3241 | 1.1011 | 0.94729 | 0.8963 | 0.87879 | 0.87331 | 0.86142 | 0.86112 | 0.86362 | 0.86704 | 0.87348 | 0.88039 | 0.88635 | 0.89143 | 0.9001 | 0.90707
|  |  |  | 160 | 1.7412 | 1.7412 | 1.6495 | 1.5578 | 1.3286 | 1.0994 | 0.93709 | 0.87994 | 0.86313 | 0.85766 | 0.8451 | 0.84489 | 0.84773 | 0.8516 | 0.85978 | 0.86682 | 0.87386 | 0.87943 | 0.8898 | 0.89799
|  |  |  | 162 | 1.5153 | 1.5153 | 1.452 | 1.3886 | 1.2303 | 1.0719 | 0.9135 | 0.86243 | 0.84614 | 0.84029 | 0.82694 | 0.82636 | 0.82999 | 0.83459 | 0.84436 | 0.8521 | 0.85958 | 0.86654 | 0.87813 | 0.88695
|  |  |  | 164 | 1.3818 | 1.3818 | 1.3308 | 1.2799 | 1.1526 | 1.0253 | 0.89805 | 0.8429 | 0.8264 | 0.82014 | 0.80558 | 0.80546 | 0.80953 | 0.81482 | 0.82544 | 0.83521 | 0.84422 | 0.85192 | 0.86551 | 0.87602
|  |  |  | 166 | 1.9612 | 1.9612 | 1.8453 | 1.7293 | 1.4394 | 1.1495 | 0.85962 | 0.81933 | 0.80421 | 0.79682 | 0.7808 | 0.78063 | 0.78595 | 0.79226 | 0.80444 | 0.81571 | 0.82609 | 0.83508 | 0.85091 | 0.86312
|  |  |  | 168 | 1.0363 | 1.0363 | 1.0161 | 0.99597 | 0.94557 | 0.89517 | 0.84477 | 0.79437 | 0.77944 | 0.77042 | 0.75322 | 0.75365 | 0.75957 | 0.76689 | 0.78155 | 0.79426 | 0.80599 | 0.81642 | 0.8346 | 0.84897
|  |  |  | 170 | 0.89409 | 0.89409 | 0.88426 | 0.87444 | 0.84988 | 0.82532 | 0.80076 | 0.7762 | 0.75164 | 0.74135 | 0.72135 | 0.72203 | 0.72865 | 0.73745 | 0.75418 | 0.76947 | 0.78336 | 0.79564 | 0.81652 | 0.83388
|  |  |  | 172 | 0.87997 | 0.87997 | 0.87004 | 0.8601 | 0.83526 | 0.81042 | 0.78558 | 0.76074 | 0.7359 | 0.71106 | 0.6855 | 0.68615 | 0.69475 | 0.70488 | 0.7246 | 0.74171 | 0.75779 | 0.77213 | 0.79624 | 0.81528
|  |  |  | 174 | 0.62623 | 0.62623 | 0.62665 | 0.62706 | 0.6281 | 0.62914 | 0.63019 | 0.63123 | 0.63227 | 0.63331 | 0.6406 | 0.6458 | 0.65621 | 0.66867 | 0.69203 | 0.71147 | 0.73011 | 0.74652 | 0.77435 | 0.7965
|  |  |  | 176 | 0.57481 | 0.57481 | 0.57538 | 0.57596 | 0.5774 | 0.57884 | 0.58028 | 0.58172 | 0.58316 | 0.5846 | 0.59468 | 0.60188 | 0.61628 | 0.63043 | 0.65825 | 0.67795 | 0.69872 | 0.71787 | 0.74907 | 0.77341
|  |  |  | 178 | 0.5361 | 0.5361 | 0.53668 | 0.53726 | 0.53872 | 0.54017 | 0.54162 | 0.54307 | 0.54452 | 0.54598 | 0.55614 | 0.5634 | 0.57792 | 0.59451 | 0.62487 | 0.63775 | 0.65971 | 0.68189 | 0.71726 | 0.74559
|  |  |  | 180 | 0.49744 | 0.49744 | 0.49803 | 0.49861 | 0.50008 | 0.50154 | 0.503 | 0.50446 | 0.50592 | 0.50739 | 0.51762 | 0.52493 | 0.53955 | 0.55858 | 0.59149 | 0.59755 | 0.62071 | 0.64591 | 0.68546 | 0.71776