## mPDR-v1
| TG-43 CONSENSUS |  |  |  |  | Nucletron PDR 192Ir mPDR-v1
Dose calculation for photon-emitting brachytherapy sources with average energy higher than 50 keV: Full Report of the AAPM and ESTRO
Med. Phys. 39 (2012) 2904-2929
| L = | 1.12 |  | cGy/(h U)
Interpolated / extrapolated data are boldface / underlined. Values inside the source are in italics.
| gL( r) |  |  |  | F(r,q) | [13 empty] | QA Along-away
| L = | 0.12 | cm |  |  |  | Distance from Active Source Center (cm)
| r (cm) | gL(r ) unbounded |  | q (º) | 0 | 0.25 | 0.5 | 1 | 2 | 3 | 4 | 5 | 7.5 | 10 | z (cm) | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7
| 0 | 1.0106 |  | 0 | 0.85369 | 0.85369 | 0.8753 | 0.8862 | 0.8834 | 0.8902 | 0.8942 | 0.8989 | 0.9162 | 0.9269 | 7 | 0.020656 | 0.020652 | 0.020698 | 0.020646 | 0.020544 | 0.020149 | 0.019562 | 0.0181 | 0.016262 | 0.0143 | 0.012363 | 0.010686
| 0.2 | 1.0106 |  | 1 | 0.85464 | 0.85464 | 0.8756 | 0.879 | 0.883 | 0.8908 | 0.8996 | 0.9021 | 0.9157 | 0.9282 | 6 | 0.028241 | 0.028306 | 0.028307 | 0.028192 | 0.027993 | 0.027198 | 0.026214 | 0.023535 | 0.020556 | 0.017628 | 0.014915 | 0.01252
| 0.25 | 1.0067 |  | 2 | 0.85563 | 0.85563 | 0.8753 | 0.8799 | 0.8845 | 0.8927 | 0.8991 | 0.9025 | 0.9163 | 0.9278 | 5 | 0.040601 | 0.040809 | 0.040691 | 0.040455 | 0.039982 | 0.038467 | 0.036531 | 0.031559 | 0.026434 | 0.021764 | 0.017851 | 0.014582
| 0.5 | 1.0021 |  | 3 | 0.85667 | 0.85667 | 0.8736 | 0.8814 | 0.8876 | 0.8967 | 0.9025 | 0.9063 | 0.9202 | 0.9314 | 4 | 0.063051 | 0.063482 | 0.0633 | 0.06262 | 0.061246 | 0.057973 | 0.053392 | 0.04353 | 0.034375 | 0.026853 | 0.021103 | 0.016828
| 0.75 | 1.001 |  | 5 | 0.85888 | 0.85888 | 0.8788 | 0.8863 | 0.8933 | 0.8994 | 0.907 | 0.9085 | 0.9239 | 0.9332 | 3 | 0.11195 | 0.11229 | 0.11119 | 0.10863 | 0.10497 | 0.094938 | 0.083366 | 0.061283 | 0.044459 | 0.032764 | 0.024605 | 0.018998
| 1 | 1 |  | 7 | 0.86127 | 0.86127 | 0.8852 | 0.895 | 0.8984 | 0.9055 | 0.9114 | 0.9124 | 0.9273 | 0.9377 | 2 | 0.24913 | 0.24956 | 0.24218 | 0.22935 | 0.2122 | 0.17388 | 0.13793 | 0.085992 | 0.056184 | 0.038719 | 0.027882 | 0.020837
| 1.5 | 1.0023 |  | 10 | 0.86661 | 0.86661 | 0.8912 | 0.8981 | 0.9037 | 0.9092 | 0.9186 | 0.9191 | 0.9322 | 0.9408 | 1.5 | 0.4427 | 0.43826 | 0.41503 | 0.37626 | 0.3312 | 0.24463 | 0.17806 | 0.10006 | 0.061766 | 0.041198 | 0.029182 | 0.021608
| 2 | 1.0051 |  | 12 | 0.87172 | 0.87172 | 0.8971 | 0.9036 | 0.9078 | 0.9145 | 0.9206 | 0.9215 | 0.9337 | 0.9443 | 1 | 0.99732 | 0.9645 | 0.84554 | 0.692 | 0.54903 | 0.34271 | 0.22442 | 0.11291 | 0.066296 | 0.043304 | 0.030305 | 0.022192
| 3 | 1.0089 |  | 15 | 0.87536 | 0.87536 | 0.9042 | 0.9143 | 0.9141 | 0.9191 | 0.9237 | 0.9244 | 0.9357 | 0.944 | 0.5 | 3.9916 | 3.3797 | 2.1913 | 1.3662 | 0.89338 | 0.44931 | 0.26492 | 0.12236 | 0.069472 | 0.044623 | 0.030933 | 0.022525
| 4 | 1.0059 |  | 20 | 0.88654 | 0.88654 | 0.9161 | 0.9238 | 0.9258 | 0.9308 | 0.9367 | 0.9369 | 0.9452 | 0.9548 | 0 | 1451140526.7598 | 17.7271 | 4.4732 | 1.9913 | 1.12 | 0.49927 | 0.28168 | 0.12569 | 0.07049 | 0.045158 | 0.031173 | 0.022629
| 5 | 1.0068 |  | 25 | 0.89864 | 0.89864 | 0.9273 | 0.9367 | 0.9359 | 0.9404 | 0.945 | 0.9429 | 0.9526 | 0.9589 | -0.5 | 3.1753 | 3.3578 | 2.1974 | 1.3718 | 0.89554 | 0.4493 | 0.26533 | 0.12235 | 0.069524 | 0.044553 | 0.030923 | 0.022544
| 6 | 1.0008 |  | 30 | 0.91056 | 0.91056 | 0.9398 | 0.9476 | 0.9475 | 0.9506 | 0.9518 | 0.9499 | 0.9596 | 0.9638 | -1 | 0.79081 | 0.904 | 0.83713 | 0.69336 | 0.55034 | 0.3431 | 0.22426 | 0.11287 | 0.066413 | 0.043324 | 0.030255 | 0.022164
| 8 | 0.98284 |  | 35 | 0.92783 | 0.92783 | 0.9501 | 0.9574 | 0.9585 | 0.9613 | 0.9632 | 0.9606 | 0.9685 | 0.9742 | -1.5 | 0.35479 | 0.39022 | 0.40076 | 0.37243 | 0.3304 | 0.24472 | 0.17828 | 0.09999 | 0.061524 | 0.041216 | 0.029272 | 0.021611
| 10 | 0.94688 |  | 40 | 0.94079 | 0.94079 | 0.962 | 0.9699 | 0.967 | 0.97 | 0.9719 | 0.9692 | 0.9747 | 0.9751 | -2 | 0.20178 | 0.21226 | 0.22842 | 0.22373 | 0.21001 | 0.17347 | 0.13817 | 0.086204 | 0.056095 | 0.038509 | 0.02783 | 0.020899
|  |  |  | 45 | 0.95204 | 0.95204 | 0.9711 | 0.9771 | 0.9763 | 0.9761 | 0.9781 | 0.9764 | 0.9829 | 0.9846 | -3 | 0.092795 | 0.093494 | 0.10031 | 0.10296 | 0.10167 | 0.094007 | 0.083151 | 0.061325 | 0.044588 | 0.032729 | 0.024555 | 0.018958
|  |  |  | 50 | 0.96425 | 0.96425 | 0.9794 | 0.9856 | 0.9836 | 0.9851 | 0.9859 | 0.983 | 0.9862 | 0.9903 | -4 | 0.05349 | 0.053665 | 0.055359 | 0.057629 | 0.058256 | 0.056482 | 0.052893 | 0.043431 | 0.034321 | 0.02687 | 0.021104 | 0.016823
|  |  |  | 55 | 0.9824 | 0.9824 | 0.9855 | 0.988 | 0.9879 | 0.9876 | 0.9895 | 0.9853 | 0.9908 | 0.9911 | -5 | 0.034661 | 0.034925 | 0.035407 | 0.036549 | 0.037255 | 0.037099 | 0.035719 | 0.031475 | 0.026362 | 0.021732 | 0.017833 | 0.014582
|  |  |  | 60 | 0.9883 | 0.9883 | 0.9887 | 0.9953 | 0.9924 | 0.9923 | 0.9937 | 0.9924 | 0.9932 | 0.9947 | -6 | 0.024438 | 0.024519 | 0.024653 | 0.025166 | 0.025746 | 0.025992 | 0.025459 | 0.023276 | 0.020473 | 0.017587 | 0.014887 | 0.012495
|  |  |  | 65 | 0.9916 | 0.9916 | 0.9927 | 0.9957 | 0.9964 | 0.997 | 0.9985 | 0.9954 | 0.9963 | 0.9988 | -7 | 0.018111 | 0.018117 | 0.018181 | 0.018384 | 0.018698 | 0.019061 | 0.018919 | 0.017786 | 0.016176 | 0.014245 | 0.012372 | 0.010661
|  |  |  | 70 | 0.9941 | 0.9941 | 0.9958 | 0.9999 | 0.9981 | 0.9984 | 1.0001 | 0.9973 | 0.9993 | 1.0006
|  |  |  | 75 | 0.9956 | 0.9956 | 0.9966 | 1.0019 | 0.9987 | 0.9983 | 0.999 | 0.9945 | 0.9976 | 0.9997
|  |  |  | 80 | 0.9982 | 0.9982 | 0.9993 | 1.008 | 1.0008 | 1.0008 | 1.0025 | 0.9989 | 1.0019 | 1.0001
|  |  |  | 85 | 0.9993 | 0.9993 | 0.9983 | 1.0061 | 0.9996 | 0.9992 | 0.9999 | 0.998 | 1.0014 | 1.0041
|  |  |  | 90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  |  | 95 | 0.9996 | 0.9996 | 0.9998 | 1.0045 | 1.0009 | 1.0019 | 1.0022 | 0.9963 | 1.0031 | 1.0043
|  |  |  | 100 | 0.9985 | 0.9985 | 0.9975 | 1.0014 | 1.0002 | 1.0004 | 1.0013 | 0.9981 | 0.9988 | 1.0024
|  |  |  | 105 | 0.9969 | 0.9969 | 0.9994 | 1.0036 | 1.0007 | 1.0011 | 1.0012 | 0.9988 | 1.0019 | 1.0047
|  |  |  | 110 | 0.9953 | 0.9953 | 0.9987 | 0.9985 | 0.9976 | 0.9968 | 0.9971 | 0.9898 | 0.9971 | 0.997
|  |  |  | 115 | 0.9927 | 0.9927 | 0.9953 | 0.9983 | 0.9956 | 0.9964 | 0.9979 | 0.992 | 0.9943 | 0.9969
|  |  |  | 120 | 0.9887 | 0.9887 | 0.991 | 0.9986 | 0.992 | 0.9907 | 0.9943 | 0.9899 | 0.9926 | 0.9957
|  |  |  | 125 | 0.9838 | 0.9838 | 0.9872 | 0.993 | 0.9884 | 0.991 | 0.9929 | 0.9888 | 0.9907 | 0.9916
|  |  |  | 130 | 1 | 1 | 0.9798 | 0.9858 | 0.9843 | 0.9843 | 0.9866 | 0.9848 | 0.9854 | 0.9878
|  |  |  | 135 | 0.969 | 0.969 | 0.973 | 0.981 | 0.9764 | 0.9782 | 0.9795 | 0.9748 | 0.9815 | 0.9822
|  |  |  | 140 | 0.9574 | 0.9574 | 0.9626 | 0.973 | 0.9678 | 0.9691 | 0.9709 | 0.9674 | 0.9718 | 0.978
|  |  |  | 145 | 0.9463 | 0.9463 | 0.951 | 0.9604 | 0.9555 | 0.957 | 0.9616 | 0.9582 | 0.9646 | 0.9697
|  |  |  | 150 | 0.934 | 0.934 | 0.9385 | 0.9475 | 0.9429 | 0.9468 | 0.9511 | 0.9486 | 0.9548 | 0.96
|  |  |  | 155 | 0.91795 | 0.91795 | 0.9197 | 0.9232 | 0.9239 | 0.928 | 0.933 | 0.93 | 0.9387 | 0.9452
|  |  |  | 160 | 0.89768 | 0.89768 | 0.89825 | 0.8994 | 0.9017 | 0.9072 | 0.9113 | 0.9108 | 0.9241 | 0.9347
|  |  |  | 165 | 0.8576 | 0.8576 | 0.8592 | 0.8624 | 0.8688 | 0.8776 | 0.8852 | 0.8871 | 0.9032 | 0.9149
|  |  |  | 168 | 0.83002 | 0.83002 | 0.83165 | 0.8349 | 0.8414 | 0.8521 | 0.8613 | 0.8632 | 0.884 | 0.8971
|  |  |  | 170 | 0.80297 | 0.80297 | 0.80475 | 0.8083 | 0.8154 | 0.8282 | 0.8385 | 0.8467 | 0.8687 | 0.8849
|  |  |  | 173 | 0.7325 | 0.7325 | 0.7367 | 0.7451 | 0.7619 | 0.7796 | 0.7952 | 0.8067 | 0.8354 | 0.8591
|  |  |  | 175 | 0.68515 | 0.68515 | 0.6916 | 0.7045 | 0.7303 | 0.7493 | 0.7695 | 0.7832 | 0.8178 | 0.8425
|  |  |  | 177 | 0.6884 | 0.6884 | 0.693 | 0.7022 | 0.7206 | 0.742 | 0.7617 | 0.7756 | 0.8099 | 0.8383
|  |  |  | 178 | 0.68567 | 0.68567 | 0.69035 | 0.6997 | 0.7184 | 0.7385 | 0.7584 | 0.7721 | 0.8088 | 0.8353
|  |  |  | 179 | 0.67845 | 0.67845 | 0.6839 | 0.6948 | 0.7166 | 0.7364 | 0.7567 | 0.7691 | 0.8074 | 0.8336
|  |  |  | 180 | 0.6931 | 0.6931 | 0.6963 | 0.7027 | 0.7155 | 0.7379 | 0.7586 | 0.7674 | 0.8085 | 0.8318